EMEP/CCC-Report 10/2003

EMEP POP laboratory comparison
2000–2002

Stein Manø and Jan Schaug

| NILU | : | EMEP/CCC-Report 10/2003 |
| --- | --- | --- |
| REFERENCE | : | O-100116 |
| DATE | : | FEBRUARY 2004 |

EMEP Co-operative Programme for Monitoring and Evaluation of the Long-range Transmission of Air Pollutants 
in Europe

EMEP POP laboratory comparison
2000–2002

Stein Manø and Jan Schaug

|  | Norwegian Institute for Air Research P.O. Box 100, N-2027 Kjeller, Norway |
| --- | --- |

Contents

Page

1.	Introduction	5

2.	Organisation of the laboratory comparison	5

3.	Preparation of samples for round 1	6

4.	Quality assurance	7

5.	Participating laboratories	8

6.	Results from round 1	9

6.1	Results for organochlorine compounds round 1	9

6.2	Results for PAH round 1	11

7.	Preparation of samples for round 2	13

8.	Results from round 2	14

8.1	Results for organochlorine compounds round 2	14

8.2	Results for PAH round 2	22

9.	Conclusion	30

10.	References	31

Annex 1 List of participants	33

Annex 2 Methods	41

Annex 3 Organochlorine compounds, round 1	65

Annex 4 PAH compounds, round 1	81

Annex 5 Organochlorine compounds, round 2, sample 1 and 2	103

Annex 6 PAH compounds, round 2, sample 1 and 2	135

EMEP POP laboratory comparison
2000–2002

# Introduction

The first collection of POP data in EMEP was reported in 1996 (Berg et al.), while the more formal inclusion in the EMEP measurement programme came in 1999. The activity on POPs had been initiated with the “Workshop on emission inventories and modelling of heavy metals and persistent organic pollutants”, in Bilthoven, Netherlands, May 1994, and the Steering Body to EMEP instructed the CCC to start the collection of POPs and heavy metal data later that year.

The “EMEP Expert meeting on measurements of POP in air and precipitation” in Lillehammer, Norway, November 1997 (Lükewille, 1998), concluded that an analytical comparison for laboratories before the (POP) network was put in operation was essential, and that a comparison should be arranged. The Steering Body agreed to this later that year and the preparations were started.

After a period with procurement of standards and other necessities and preparations, the first round of samples were distributed in November 2000 while the second round with samples were sent to the participants in August 2001.

# Organisation of the laboratory comparison

27 laboratories responded to the invitation and indicated that they would take part in the exercise. A few laboratories asked for more than one set of samples to be distributed to different laboratories.

Table 1 shows the 19 names of the polycyclic aromatic hydrocarbons (PAH) and the 14 organochlorine compounds included in this comparison.

The comparison was carried out as a two-step exercise. During the first round each laboratory received two ampoules with the chlorinated compounds and two ampoules with a PAH mixture. One of the two ampoules from each group of components contained standards with known concentrations; a list of components in these solutions was distributed separately. The other two ampoules contained the same components but at concentrations unknown to the laboratories. The upper and lower limits of the concentrations in these unknown mixtures were also given to the laboratories.

This first round should therefore be a pure analytical exercise with no need for pre-concentrating or clean-up. The laboratories were asked to analyse the two unknown mixtures using their own standards in triplicate and to forward all the results to the CCC. The known standards were intended for a check on the laboratories own standards, but these results should not be sent the CCC. The laboratories were also asked to give a brief description of their analytical methods (Annex 2) with estimates of detection limits and uncertainties.

Table 1:	PAH and organochlorine compounds measured in the laboratory comparison.

| PAHs | Organochlorine compound |
| --- | --- |
| Naphtalene | p,p’-DDT |
| Acenaphthylene | p,p’-DDE |
| Acenaphthene | -Chlordane |
| Fluorene | -Chlordane |
| Phenanthrene | Hexachlorocyclohexanes, HCH: |
| Anthracene | -HCH |
| Fluoranthene | -HCH |
| Pyrene | Hexachlorobenzene, HCB |
| Benzo[a]anthracene | Polychlorinated biphenyls: |
| Chrysene | PCB 28 |
| Benzo[b]fluoranthene | PCB 52 |
| Benzo[k]fluoranthene | PCB 101 |
| Benzo[a]pyrene | PCB 118 |
| Indeno[1,2,3-cd]pyrene | PCB 138 |
| Dibenzo[a,h]anthracene | PCB 153 |
| Benzo[ghi]perylene | PCB 180 |
| Biphenyl |  |
| Perylene |  |
| Benzo[e]pyrene |  |

# Preparation of samples for round 1

Pre-experiments with flame sealing of ampoules showed a large tendency to sooth formation below the sealing point, inside the ampoule, if it contained the solvents hexane or cyclohexane. By repeating the test with ampoules containing the less volatile solvent isooctane sooth formation was avoided. Hence it was decided to use isooctane as solvent in round 1. This, however, caused a difficulty for laboratories using HPLC for the determination of PAHs, where polar solvents like acetonitrile or methanol are preferred, since a solvent change had to be made.

Solid standard compounds for all compounds but perylene were dissolved in isooctane to produce primary standard solutions. For perylene, which was difficult to dissolve in isooctane, a mixture of isooctane and toluene was used.

From the primary homemade standards and the purchased liquid standards, volumes were taken out using syringes to produce a secondary standard. The actual amount taken out from each primary standard to the secondary standard was determined by weighing each time, before the final dilution with isooctane.

The ampoules were filled and flame sealed using a propane gas burner.

The ampoules were packed in sealed plastic boxes filled with shock absorbing plastic foil. A test was performed where a box containing filled ampoules was dropped from the third floor onto a concrete surface on the ground without causing any damage on the ampoules.

Finally a transport agency packed the boxes with ampoules in accordance with the safety regulations for transport of inflammable liquids with air cargo before the shipment.

In spite of shipping the samples in accordance to all rules (shipping and custom) several participants had problems with both shipping companies and custom authorities.

# Quality assurance

Table 2 shows the purity of the standard compounds used in round 1 and the name of the company, which supplied them.

Table 2:	Standard compounds used for round 1 and their purity.

| Compound | Purity % | Solution/solid standard | Supplier |
| --- | --- | --- | --- |
| Naphthalene | 99 | Solution with certificate | Restek |
| Acenaphthylene | 99 | Solution with certificate | Restek |
| Acenaphthene | 99 | Solution with certificate | Restek |
| Fluorene | 99 | Solution with certificate | Restek |
| Phenanthrene | 99 | Solution with certificate | Restek |
| Anthracene | 99 | Solution with certificate | Restek |
| Fluoranthene | 98 | Solution with certificate | Restek |
| Pyrene | 99 | Solution with certificate | Restek |
| Benzo[a]anthracene | 99 | Solution with certificate | Restek |
| Chrysene | 99 | Solution with certificate | Restek |
| Benzo[b]fluoranthene | 99 | Solution with certificate | Restek |
| Benzo[k]fluoranthene | 99 | Solution with certificate | Restek |
| Benzo[a]pyrene | 99 | Solution with certificate | Restek |
| Indeno[1,2,3-cd]pyrene | 99 | Solution with certificate | Restek |
| Dibenzo[a,h]anthracene | 99 | Solution with certificate | Restek |
| Benzo[ghi]perylene | 99 | Solution with certificate | Restek |
| Biphenyl | >99 | Solid* | AccuStandard |
| Perylene | >99 | Solid* | AccuStandard |
| Benzo[e]pyrene | >99 | Solid* | AccuStandard |
| p,p’-DDT | 99 | Solution with certificate | Restek |
| p,p’-DDE | 99 | Solution with certificate | Restek |
| γ -chlordane | 99 | Solution with certificate | Restek |
| -chlordane | 99 | Solution with certificate | Restek |
| γ-HCH | 99 | Solution with certificate | Restek |
| α-HCH | 99 | Solution with certificate | Restek |
| Hexachlorobenzene, HCB | 99 | Solution with certificate | Restek |
| PCB 28 | >99 | Solid* | Promochem |
| PCB 52 | >99 | Solid* | Promochem |
| PCB 101 | >99 | Solid* | Promochem |
| PCB 118 | >99 | Solid* | Promochem |
| PCB 138 | >99 | Solid* | Promochem |
| PCB 153 | >99 | Solid* | Promochem |
| PCB 180 | >99 | Solid* | Promochem |

* All solid standards were controlled on purity by NILU using GC/MS in scan and EI mode in mass range: 50-550. The standard compound peak area was 99% of total ion chromatogram area or better.

1st round:

- The ampoules were weighed after sealing and weighed again one week later to detect leaking ampoules.

- Some ampoules were analysed at the CCC using GC/MS and isotopically labelled internal standards before the distribution to the laboratories.

2nd round:

- The ampoules were weighed after sealing and weighed again one week later to detect leaking ampoules.

# Participating laboratories

Most of the laboratories took part in both round 1 and 2 and measured both PAH and organochlorine compounds. Some laboratories returned results from parts of the exercise only and a few laboratories participated only in one of the rounds. The 21 laboratories which returned results are listed in Table 3. More information is given in Annex 1.

Table 3:	Laboratories participating in the POP laboratory comparison.

| Country | Participating institution |
| --- | --- |
| Austria | Umweltbundesamt, Vienna |
| Belgium | Vlaamse Milieumaatschapij, Gent |
| Canada | Meteorological Service of Canada, Downsview |
| Canada | Freshwater Institute, DFO, Winnipeg |
| Czech Republic | RECETOX, Brno |
| France | Lab Nat. d'Essais, Paris |
| France | Lab. Wolff Environ., Evry |
| France | Calydra, Paimboeuf |
| France | Lab. LARA, Toulouse |
| Germany | UBA Labor für Wasseranalytik, Berlin |
| Germany | Alfred-Wegener-Inst., Bremerhaven |
| Germany | IfE Analytik GmbH, Leipzig |
| Iceland | Dep. Pharm. & Tox., Univ. Iceland, Reykjavik |
| Ireland | Environmental Pollution Agency, Dublin |
| Lithuania | Joint Res. Centre, Vilnius |
| Lithuania | Institute of Physics, Vilnius |
| Republic of Macedonia | Republic Hydrometeorological Institute, Skopje |
| Norway | Norwegian Institute for Air Research, Kjeller |
| United Kingdom | AEA Technology, Culham |
| United Kingdom | University of Lancaster, Lancaster |
| Yugoslavia | Federal Hydrometeorological Institute, Belgrade |

# Results from round 1

The results are given below in Table 4 – Table 7 and further presented in 
Annex 3–4. The first round of the comparison gave results that are related to the calibration and the instrumental performance.

Firstly, for all compounds an average value and a standard deviation was calculated from the data sent from all laboratories.

Secondly a range from [average-(2*st.dev.)] to [average+(2*st.dev.)] was computed. Finally a new average value was computed for each compound, excluding all reported values above or below the range computed in step two. This is the average value reported in Annex 3 and 4.

## Results for organochlorine compounds round 1

Simple statistics and the ratios between arithmetic means and medians to the expected concentrations are given in Table 4 below.

Table 4:	Statistical overview for organochlorine compounds in round 1. Concentration unit: pg/l.

| Organochlorine compounds | Average | Median | Expected | St.dev. | Average/ Expected | Median/ Expected |
| --- | --- | --- | --- | --- | --- | --- |
| p,p'-DDT | 21.7 | 21.9 | 20.0 | 6.0 | 1.08 | 1.09 |
| p,p'-DDE | 40.9 | 40.1 | 40.0 | 10.8 | 1.02 | 1.00 |
| -Chlordane | 29.5 | 10.7 | 7.0 | 57.9 | 4.21 | 1.52 |
| -Chlordane | 15.1 | 14.9 | 15.0 | 3.9 | 1.00 | 0.99 |
| -HCH | 275.5 | 284.3 | 300.0 | 53.5 | 0.92 | 0.95 |
| -HCH | 155.9 | 114.7 | 120.0 | 180.6 | 1.30 | 0.96 |
| HCB | 139.3 | 141.5 | 150.0 | 19.6 | 0.93 | 0.94 |
| PCB 28 | 37.9 | 37.9 | 36.0 | 7.8 | 1.05 | 1.05 |
| PCB 52 | 70.0 | 68.5 | 76.0 | 5.9 | 0.92 | 0.90 |
| PCB 101 | 35.3 | 38.2 | 42.0 | 13.8 | 0.84 | 0.91 |
| PCB 118 | 56.1 | 53.0 | 53.0 | 17.1 | 1.06 | 1.00 |
| PCB 138 | 22.9 | 24.6 | 24.0 | 7.3 | 0.95 | 1.02 |
| PCB 153 | 21.7 | 21.2 | 22.0 | 4.5 | 0.99 | 0.96 |
| PCB 180 | 28.5 | 29.5 | 30.0 | 4.7 | 0.95 | 0.98 |

Inspections of the central tendencies in the group of laboratories show that the ratio median to expected is good and vary from 1.10 to 0.90 for the pesticides with the γ -chlordane results as the one exception with a ratio at 1.52. The results for the organochlorine compounds had large outliers particularly for p,p’-DDT, p,p’-DDE, γ-chlordane, and for several PCBs. This is seen from the standard deviation and the ratios average to expected. γ-chlordane, present at low concentration, was the most challenging compound of the organochlorine group to quantify.

The averages of the three single measurements reported from each laboratory are given in Table 5 below and in Annex 3.

Table 5:	Round 1; results of the analyses of organochlorine compounds in a mixture of standards.  The expected concentrations are given in the lower rightmost column. Unit: pg/μl.

| Lab. Code | 101 | 102 | 103 | 104 | 105 | 108 | 109 | 110 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| p,p'-DDT | 19.6 | 21.8 | 20.2 | 26.7 | 34.4 | 22.3 | 6.8 | 24.0 |
| p,p'-DDE | 39.0 | 47.3 | 40.1 | 47.9 | 66.4 | 34.7 | 18.7 |  |
| -Chlordane | 9.8 | 11.0 | 10.7 | 11.7 | 10.6 | 7.3 | 34.7 |  |
| -Chlordane | 14.6 | 19.7 | 17.3 | 17.9 | 13.3 | 11.0 | 15.0 |  |
| -HCH | 328 | 325 | 295 | 272 | 104 | 279 | 274 | 292 |
| -HCH | 118 | 130 | 104 | 112 | 34 | 111 | 104 | 119 |
| HCB | 137 | 156 | 151 | 130 | 91 | 155 | 144 | 141 |
| PCB 28 | 36.3 | 43.1 | 22.6 | 39.6 | 45.3 | 38.4 | 31.0 | 33.7 |
| PCB 52 | 71.8 | 77.3 | 65.7 | 60.7 | 68.6 | 68.3 | 64.0 | 66.0 |
| PCB 101 | 39.6 | 38.6 | 29.0 | 37.8 | 56.6 | 43.6 | 45.7 | 37.3 |
| PCB 118 | 51.3 | 64.8 | 34.1 | 53.5 | 108.5 | 57.1 | 65.7 | 50.7 |
| PCB 138 | 21.9 | 28.4 | 11.4 | 24.8 | 36.8 | 25.2 | 27.0 | 24.3 |
| PCB 153 | 20.4 | 27.5 | 15.8 | 21.4 | 32.1 | 23.2 | 17.7 | 21.0 |
| PCB 180 | 29.6 | 33.1 | 17.9 | 28.6 | 35.9 | 29.4 | 28.7 | 27.7 |
| Lab. code | 113 | 118 | 121 | 123 | 129A | 129L | Expected |  |
| p,p'-DDT | 19.7 | 18.3 | 27.0 | 24.7 | 22.0 | 15.9 | 20.0 |  |
| p,p'-DDE | 40.5 | 38.7 | 28.0 | 48.4 | 38.0 | 44.4 | 40.0 |  |
| -Chlordane |  | 8.1 | 8.5 | 220.0 | 13.0 | 8.7 | 7.0 |  |
| -Chlordane |  | 13.9 | 6.0 | 21.4 | 16.0 | 14.8 | 15.0 |  |
| -HCH | 284 | 291 | 275 | 285 | 230 | 323 | 300.0 |  |
| -HCH | 117 | 111 | 58 | 128 | 800 | 138 | 120.0 |  |
| HCB | 128 | 155 | 142 | 124 | 120 | 175 | 150.0 |  |
| PCB 28 | 55.5 | 34.0 | 37.5 | 38.2 | 29.0 | 46.2 | 36.0 |  |
| PCB 52 | 77.4 | 78.9 | 68.0 | 71.3 | 63.0 | 78.9 | 76.0 |  |
| PCB 101 | 39.3 | 37.8 | 7.6 | 2.8 | 33.0 | 45.0 | 42.0 |  |
| PCB 118 | 49.4 | 53.9 | 36.7 | 52.4 | 44.0 | 62.9 | 53.0 |  |
| PCB 138 | 22.9 | 22.3 | 6.0 | 26.5 | 16.0 | 27.2 | 24.0 |  |
| PCB 153 | 21.8 | 20.5 | 13.9 | 25.7 | 20.0 | 23.1 | 22.0 |  |
| PCB 180 | 30.3 | 30.6 | 19.7 | 33.4 | 25.0 | 29.6 | 30.0 |  |

The results for HCB, PCB-52, and γ-HCH have the lowest number of results outside 25% from the expected, and the largest numbers of results better than 15% and better than 5% (Table 4, and Annex 3). This may be due to the fact that these compounds are among those with the highest concentrations in the exercise and that the laboratories possess experience with these substances at these concentration levels. Among the compounds with high numbers of results deviating more than 25% from the expected and relatively few good results are γ-chlordane, p,p’-DDT, p,p’– DDE, and PCB-28. These are among the compounds with the lowest concentrations in the first round of the exercise.

The laboratories performance will presumable depend on many factors, one very important being the experience and routines they have, or lack, with respect to measurements at low concentration levels as in this comparison. As will be seen below, this is even more important in the second part of the comparison than in the first round.

The laboratories with the largest number of result better than 5% from the expected in the first part of the comparison were 101, 113, and 118, followed by 108 and 110.

## Results for PAH round 1

Simple statistics and ratios between arithmetic means and median to the expected concentrations have been given in Table 6 for the PAH.

Table 6:	Statistical overview for PAH in round 1.
Concentration unit: ng/l.

| PAH | Average | Median | Expected | St.dev. | Average/ Expected | Median/ Expected |
| --- | --- | --- | --- | --- | --- | --- |
| Naphthalene | 7.73 | 8.21 | 8.33 | 2.40 | 0.93 | 0.99 |
| Acenaphthylene | 2.58 | 2.73 | 2.00 | 0.85 | 1.29 | 1.37 |
| Acenaphthene | 3.80 | 4.00 | 4.17 | 1.31 | 0.91 | 0.96 |
| Fluorene | 3.63 | 3.75 | 4.00 | 0.94 | 0.91 | 0.94 |
| Phenanthrene | 7.72 | 8.01 | 8.33 | 1.76 | 0.93 | 0.96 |
| Anthracene | 0.38 | 0.37 | 0.42 | 0.09 | 0.91 | 0.87 |
| Fluoranthene | 3.79 | 3.89 | 4.17 | 0.74 | 0.91 | 0.93 |
| Pyrene | 3.02 | 3.38 | 3.33 | 0.78 | 0.91 | 1.01 |
| Benzo[a]anthracene | 0.66 | 0.54 | 0.58 | 0.42 | 1.13 | 0.93 |
| Chrysene | 2.06 | 1.39 | 1.67 | 2.97 | 1.23 | 0.83 |
| Benzo[b]fluoranthene | 1.98 | 2.09 | 2.18 | 0.46 | 0.91 | 0.96 |
| Benzo[k]fluoranthene | 0.56 | 0.49 | 0.50 | 0.23 | 1.11 | 0.98 |
| Benzo[a]pyrene | 0.44 | 0.35 | 0.42 | 0.23 | 1.06 | 0.84 |
| Indeno[1,2,3-cd]pyrene | 1.18 | 1.20 | 1.33 | 0.17 | 0.89 | 0.90 |
| Dibenzo[a,h]anthracene | 0.22 | 0.16 | 0.17 | 0.17 | 1.28 | 0.94 |
| Benzo[ghi]perylene | 1.48 | 0.95 | 1.00 | 1.94 | 1.48 | 0.95 |
| Biphenyl | 6.63 | 7.10 | 7.98 | 2.58 | 0.83 | 0.89 |
| Perylene | 1.12 | 1.10 | 1.20 | 0.25 | 0.93 | 0.92 |
| Benzo[e]pyrene | 0.53 | 0.46 | 0.43 | 0.26 | 1.23 | 1.07 |

Most of the medians in Table 6 are less than ten per cent from the expected, with the largest deviation seen for acenaphthylene. The three compounds at the end of Table 6 were analysed by 8–9 participants only.

Table 7 show the results of the PAH analyses from all laboratories for all compounds during round 1.

Table 7:	Round 1; results of the analyses of PAH in a mixture of standards. The expected concentrations are given in the lower rightmost column. Unit: ng/l.
(1): GC/MS, (2): HPLC, Exp.: Expected concentrations.

| Lab. code | 101 | 102 | 103 | 104 | 105(1) | 105(2) | 107 | 108 | 109 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Naphthalene | 10.33 | 1.58 |  | 5.35 | 10.61 | 8.55 | 8.190 | 7.99 | 8.53 |
| Acenaphthylene | 3.19 | 0.22 | 2.48 | 1.68 | 3.57 | 2.99 | 2.680 | 2.60 | 2.97 |
| Acenaphthene | 4.58 | 0.24 |  | 2.53 | 5.15 | 4.10 | 4.045 | 3.55 | 4.07 |
| Fluorene | 4.17 | 1.30 | 3.02 | 2.16 | 4.77 | 4.13 | 3.670 | 3.58 | 3.80 |
| Phenanthrene | 9.35 | 5.28 | 7.51 | 4.98 | 10.60 | 8.68 | 7.890 | 8.01 | 8.40 |
| Anthracene | 0.520 | 0.274 | 0.314 | 0.190 | 0.510 | 0.380 | 0.336 | 0.354 | 0.353 |
| Fluoranthene | 4.24 | 2.70 | 3.52 | 2.37 | 5.02 | 4.16 | 3.855 | 3.89 | 4.07 |
| Pyrene | 3.59 | 2.32 | 3.21 | 1.88 | 4.13 | 3.45 | 3.070 | 3.35 | 3.43 |
| Benzo[a]anthracene | 0.423 | 0.516 | 0.535 | 0.267 | 0.650 | 0.630 | 0.531 | 0.396 | 0.547 |
| Chrysene | 1.13 | 1.70 | 1.28 | 0.55 | 1.47 | 1.66 | 1.575 | 1.17 | 1.53 |
| Benzo[b]fluoranthene | 1.75 | 2.22 | 2.79 | 1.10 | 2.17 | 2.08 | 1.850 |  | 2.30 |
| Benzo[k]fluoranthene | 0.513 | 0.462 | 0.476 | 0.417 | 0.620 | 0.520 | 0.400 | 0.442 | 0.530 |
| Benzo[a]pyrene | 0.330 | 0.314 | 0.347 | 0.267 | 0.540 | 0.420 | 0.305 | 0.338 | 0.393 |
| Indeno[1,2,3-cd]pyrene | 0.89 | 1.24 | 1.32 | 0.84 | 1.15 | 1.20 | 1.210 | 1.15 | 1.33 |
| Dibenzo[a,h]anthracene | 0.117 | 0.099 | 0.193 | 0.100 | 0.460 | 0.260 | 0.125 | 0.152 | 0.177 |
| Benzo[ghi]perylene | 0.89 | 1.10 | 0.81 | 0.68 | 0.95 | 1.07 | 0.935 | 0.92 | 1.10 |
| Biphenyl |  |  |  | 5.21 | 0.70 | 6.94 | 6.820 |  | 7.27 |
| Perylene |  |  |  | 0.82 | 0.82 | 1.17 | 1.305 |  | 1.43 |
| Benzo[e]pyrene |  |  | 0.806 | 0.333 | 0.460 | 0.490 | 0.380 |  | 0.317 |
| Lab. code | 110 | 122 | 123 | 129A | 129L(1) | 129L(2) | 130 | Exp. |  |
| Naphthalene |  |  | 8.16 | 8.25 | 7.52 | 5.82 |  | 8.33 |  |
| Acenaphthylene |  |  | 2.58 | 2.86 | 2.49 |  |  | 2.00 |  |
| Acenaphthene |  |  | 5.32 | 3.93 | 3.57 | 3.57 |  | 4.17 |  |
| Fluorene |  |  | 4.86 | 3.69 | 3.23 | 3.59 | 4.10 | 4.00 |  |
| Phenanthrene |  | 5.70 | 9.46 | 8.41 | 7.15 | 7.85 | 4.75 | 8.33 |  |
| Anthracene | 0.320 |  | 0.464 | 0.520 | 0.403 | 0.348 |  | 0.420 |  |
| Fluoranthene | 3.70 |  | 4.33 | 3.88 | 3.50 | 4.23 | 2.55 | 4.17 |  |
| Pyrene | 3.41 | 1.20 | 3.40 | 3.24 | 2.72 | 3.49 | 1.80 | 3.33 |  |
| Benzo[a]anthracene | 0.557 | 2.000 | 1.377 | 0.503 | 0.517 | 0.576 | 0.350 | 0.580 |  |
| Chrysene | 1.43 | 13.50 | 0.71 | 1.35 | 1.20 | 1.43 | 0.95 | 1.67 |  |
| Benzo[b]fluoranthene | 2.09 |  | 1.00 | 2.02 | 1.79 | 2.17 |  | 2.18 |  |
| Benzo[k]fluoranthene | 0.473 |  | 1.377 | 0.453 | 0.530 | 0.492 |  | 0.500 |  |
| Benzo[a]pyrene | 0.400 | 0.990 | 1.079 | 0.350 | 0.353 | 0.356 | 0.250 | 0.420 |  |
| Indeno[1,2,3-cd]pyrene | 1.23 |  | 1.35 | 1.11 | 1.05 |  |  | 1.33 |  |
| Dibenzo[a,h]anthracene | 0.147 | 0.083 | 0.742 | 0.163 | 0.263 | 0.160 |  | 0.170 |  |
| Benzo[ghi]perylene | 0.94 | 8.70 | 1.24 | 0.86 | 0.79 | 0.97 |  | 1.00 |  |
| Biphenyl |  |  | 9.78 | 8.26 | 6.58 |  |  | 7.98 |  |
| Perylene |  |  | 0.93 | 1.19 | 1.01 |  | 1.10 | 1.20 |  |
| Benzo[e]pyrene |  |  | 1.156 | 0.377 | 0.360 |  |  | 0.430 |  |

Benzo[k]fluoranthene, benzo[b]fluoranthene, pyrene, and naphthalene have been best measured by the laboratories when judged from the highest number of results not more than 10% from the expected, i.e. 50–70% of the results were very good. These compounds each had not more than one result 50% or higher from the expected concentration (Annex 4). The two compounds which the group of laboratories have analysed with the lowest accuracy during the first round of the comparison were acenaphthylene and dibenzo[a,h]anthracene.

The results from the first round indicated that the group of participating laboratories had a better performance for the analysis of the PAH than for the organochlorine compounds. The PAH results have less large outliers than the pesticides with a smaller range for the ratio average to expected.

The laboratory results in Table 7 reveal large performance differences among the participants. Five laboratories were able to measure more than half of the individual PAHs better than 10% from the expected value; Lab. 105, 107, 109 with HPLC, 110, 129A, and 129L with HPLC, while two laboratories failed to obtain any results within 10% (Annex 4).

# Preparation of samples for round 2

The second round in the exercise involved analysis of real sample extracts. The samples analysed were pooled extracts of high volume ambient air samples (glass fibre filters and PUF-plugs) collected at NILU during the period from 13.12.99 to 18.4.01. The filters and PUF adsorbents were extracted in four batches. Two batches, intended for determination of chlorinated compounds, were extracted with hexane (Cl-POP1 and Cl-POP2). The other two were extracted with cyclohexane (PAH1 and PAH2). For each extraction about 2.5 litres of solvent was used. The raw-extract was pre-concentrated to about 50 ml using a TurboVap 500. The extracts contained some dust and humidity at this stage. Each extract was passed trough a column packed with 5 gram of sodium sulphate (pre-flushed with 30 ml of solvent). The dried, dust free sample was collected in an Erlenmeyer flask, which was weighed before (empty) and after collection of sample. The amount of solvent was adjusted, by weight, to about 60 ml. The dried sample extract was filled into 60 single 1 ml glass ampoules, which were sealed. This was repeated for all four pooled extracts. With this volume, each ml would be equivalent to an air sample volume of about 500 m3, ~ 900 m3 for Cl-POP2.

The actual sample volumes were:

Cl-POP1:	30.148 m3, sampling during May 2000

Cl-POP2:	53.417 m3, sampling during February, March and April 2001

PAH1:	30.021 m3, sampling during December 1999, January, February, March 2000

PAH2:	38.804 m3, sampling March, April 2001

Since several samples were taken over time intervals much longer than 24 hours and losses by evaporation through the sampler are expected to occur, the analysis of the extract would not give a correct concentration of POPs in the air sampled. Sampling provided, however, realistic real life sample extracts from air samples containing POPs and possible interferences from the sampling materials and/or the air sampled. The purpose of round 2 was to test the laboratories ability to deal with such samples in a quantitative best way.

The problem with sooth formation inside the ampoules mentioned in Chapter 3 was avoided in the following way:

The 1 ml sub-sample for the ampoule was taken out from the Erlenmeyer flask by using a syringe. The vial was filled with nitrogen gas by holding a micropipette connected to a nitrogen gas bottle with a Teflon tube inside the ampoule for some seconds. Then the sample was transferred from the syringe to the ampoule and the lower part of the ampoule was immersed in liquid argon, freezing out the solvent. After freezing out the solvent, the neck of the ampoule was quickly flame sealed using a propane burner.

The ampoules were packed and transported as described for round 1.

In spite of shipping the samples in accordance to all rules (shipping and customs) several participants had problems with both shipping companies and customs authorities. One shipment was returned to CCC from the customs authorities abroad and one shipment disappeared and had to be replaced with a new one, causing delays in the project.

A set of four ampoules was sent to each participant. The samples should be handled as normal air samples for PAH and for organochlorine compounds and had to be analysed through the normal laboratory routines including the quality assurance steps. The second round therefore included the laboratory routines for handling the samples as well as the chemical analysis.

For simplicity an approximate total sampling volume of 30.000 m3 was assigned to each large sample extract and each lab was asked to report the concentration of the compounds determined relative to a sample volume of 500 m3.

# Results from round 2

The results are given below in Table 8 and Table 11 and further presented in Annex 5–6. The second round of the comparison gave results that throw light on sample handling and clean-up as well as the chemical analysis with possible interferences from other compounds.

Please note that the concentrations in the Annexes are given as reported and do not express the measurement accuracy.

Some laboratories participated in the comparison’s second round only.

Internal standards may be used to compensate for changes in detector response during quantification. If added before extraction they will also compensate for sample losses during extraction and clean-up. Otherwise a loss will produce too low results. Determination of internal standard recovery gives information about the applicability of the method. No such data has been reported during this exercise.

## Results for organochlorine compounds round 2

The medians of all results above the detection limit in the second part of the comparison have been given in Table 8. In addition to the laboratories that quantified the concentrations and contributed to the median, there were a variable number of laboratories which reported the concentrations to be lower than their detection limits. This is quite understandable since the concentrations of some of the compounds in the real samples were quite low. Table 8 shows that the concentrations of some compounds were very low; γ-chlordane, p,p’-DDT, and PCB-180 all had one or both concentrations lower than 1 pg/m3.

Table 8:	Medians, and numbers of laboratories giving results above their detection limits. Medians in pg/m3.

| Compound | 1st sample |  | 2ndsample |  |
| --- | --- | --- | --- | --- |
|  | Median | Number of laboratories | Median | Number of laboratories |
| p,p'-DDT | 0.9 | 8 | 1.2 | 9 |
| p,p'-DDE | 1.7 | 9 | 2.1 | 10 |
|  -Chlordane | 0.7 | 8 | 0.9 | 7 |
| -Chlordane | 1.3 | 6 | 1.1 | 8 |
|  -HCH | 25.0 | 13 | 16.4 | 14 |
| -HCH | 13.3 | 13 | 22.5 | 14 |
| HCB | 28.9 | 15 | 62.1 | 15 |
| PCB 28 | 7.6 | 11 | 7.3 | 11 |
| PCB 52 | 6.7 | 11 | 5.4 | 11 |
| PCB 101 | 4.1 | 11 | 2.8 | 10 |
| PCB 118 | 1.7 | 10 | 1.2 | 11 |
| PCB 138 | 2.2 | 11 | 1.2 | 10 |
| PCB 153 | 2.8 | 10 | 1.8 | 11 |
| PCB 180 | 1.2 | 9 | 0.7 | 9 |

Table 9 shows concentrations of the organochlorine compounds in the sample in round 1 and in the two samples in round 2. The values for round 2 are calculated from the median for each compound, assuming a 500 m3 sample volume and a final sample extract volume of 100 l after clean-up and pre-concentration.

Table 9:	Comparison of organochlorine compound concentrations in the standard analysed in round 1 and in the samples of round 2.

| Compound | Round1 Theoretical pg/ul | Round2 Cl-POP1 median pg/ul | Round2 Cl-POP2 median pg/ul |
| --- | --- | --- | --- |
| p,p'-DDT | 20.0 | 4.5 | 5.9 |
| p,p'-DDE | 40.0 | 8.5 | 10.7 |
| -Chlordane | 7.0 | 3.7 | 4.6 |
| -Chlordane | 15.0 | 6.5 | 5.7 |
| -HCH | 300.0 | 125.0 | 82.0 |
| -HCH | 120.0 | 66.3 | 112.5 |
| HCB | 150.0 | 144.5 | 310.5 |
| PCB 28 | 36.0 | 37.8 | 36.7 |
| PCB 52 | 76.0 | 33.7 | 27.2 |
| PCB 101 | 42.0 | 20.5 | 13.8 |
| PCB 118 | 53.0 | 8.5 | 6.0 |
| PCB 138 | 24.0 | 11.2 | 6.2 |
| PCB 153 | 22.0 | 13.8 | 9.2 |
| PCB 180 | 30.0 | 6.1 | 3.4 |

The results from the second round have been evaluated in view of the deviations from the medians of all measurements as the best estimates for the air concentrations. Annex 5 presents the results from the second part in Figures as illustrated in Figure 1 and Figure 2 for α-HCH.

Table 10 gives the air concentrations measured at two sites in year 2000 (Berg et al., 2002) that can be compared with the exercise medians of the organochlorine compounds in Table 8.

Table 10:	Air concentrations of POP at sites in the Czech Republic and Sweden in 2000. Results from weekly and monthly samples. 
Pesticides and PCB units pg/m3.

| Organochlorine compounds | Košetice CZ0003R |  |  | Rörvik SE0002R |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Min | 50% | Max | Min | 50% | Max |
|  | Weekly samples |  |  | Monthly samples |  |  |
|  | pg/m3 |  |  | pg/m3 |  |  |
| p,p’-DDT | 0.50 | 3 | 10 | 0.05 | 0.80 | 3.85 |
| p,p-DDE | 4 | 25 | 75 | 0.64 | 2.27 | 16.71 |
| -HCH | 7 | 41 | 199 | 5 | 17 | 69 |
| α-HCH | 4 | 20 | 70 | 4 | 10 | 24 |
| PCB-28 | 16 | 28 | 52 | 0.96 | 2.01 | 5.98 |
| PCB-52 | 17 | 31 | 106 | 1.29 | 2.34 | 11.2 |
| PCB-101 | 13 | 22 | 48 | 1.06 | 2.35 | 7.95 |
| PCB-118 | 3 | 5 | 8 | 0.36 | 0.83 | 2.40 |
| PCB-138 | 5 | 14 | 28 | 0.62 | 1.20 | 3.28 |
| PCB-153 | 13 | 21 | 32 | 0.69 | 1.36 | 3.78 |
| PCB-180 | 3 | 4 | 17 | 0.24 | 0.46 | 1.23 |

Comments to the results for each compound are given below.

Figure 1:	Results for -HCH measurements in extract 1. Concentrations in pg/m3.

Figure 2:	Results for -HCH measurements in extract 1. Deviation in per cent from the median concentration.

p,p’-DDT

Fifteen laboratories participated, but six laboratories reported the concentrations to be lower than their detection limits. The DDT concentrations in the samples were about 1 pg/m3, a rather typical regional concentration for northern Europe (Table 10), but in the lower end of the central European regional concentrations. The best results were obtained by laboratory 103, which also had good results in the first round of the comparison. Laboratories 101 and 113 obtained excellent results in the first part, but laboratory 101 could not detect the component and laboratory 113 obtained too high results. Laboratories 129A and 121 obtained too high results compared to their good results in round 1.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 103, 118 and 131.

Cl-POP 2: Lab. 103 and 113.

p,p’-DDE

Fifteen laboratories took part in the analyses and 6 and 5 laboratories reported the concentration to be lower than their detection limit. The results were similar to the p,p’-DDT results with respect to laboratories with too high detection limits. Laboratories 118 and 103 obtained both results less than 10% from the medians as they did in the first part of the comparison. Laboratories 115 and 131 had results < 20% from the median, none of these laboratories took part in the first part of the comparison. One laboratory reported one result outside 100 % and the other result more than 50% from the median. The concentrations were similar to the p,p’-DDT concentrations.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 103, 104, 105, 109, 115, 118, and 131.

Cl-POP 2: Lab. 103, 115, 118, 123 and 131.

γ-Chlordane (trans-Chlordane)

Fourteen laboratories participated, but only 8 and 7 laboratories respectively reported concentrations higher than the detection limits. The concentrations were less than 1 pg/m3 and four laboratories reported the result as lower than their detection limit, additionally two laboratories reported the compound as not detected. Laboratories 118 and 131 had results near the median, and 115 also had results not more than 20% from the median. Two laboratories reported results more than 100% from the median, and two other laboratories had results deviating more than 50%.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 115, 118, and 131.

Cl-POP 2: Lab. 103, 109, 115, 118 and 131.

α-Chlordane (cis-Chlordane)

The number of participants was similar as for γ-chlordane, and 6 and 8 laboratories respectively were able to quantify the concentrations of this compound in the two samples. Laboratories 115, 104, and 103 had both results closer than 20% from the median. The median concentrations were low, 1-2 pg/m3. Laboratory 105 quantified -chlordane close to the median in one of the two samples but not in the other although the medians for the two samples were quite similar. Five laboratories had too high detection limits, while one another laboratory did not detect the substance, detection limit not given. One laboratory had a result more than 100%, and another laboratory more than 50%, from the medians.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 103, 104, 115, 118 and 129A.

Cl-POP 2: Lab. 103, 104, 105, 115 and 131.

γ -HCH

Sixteen laboratories measured γ–HCH and thirteen participants had two results above the detection limit. The concentrations were much higher than for the chlordanes; 15–25 pg/m3, which could be typical for regional European concentrations (Table 10). Two laboratories reported the results wrongly lower than their detection limit; one being laboratory 110 that had very good results in the first part of the comparison. The laboratories that were close to the median had obtained very good results in the first part of the exercise as well; laboratories 101, 118, and 102 had both results less than 10% from the median. Laboratories 103 and 131 obtained results less than 20% away. Six laboratories had one or more result outside 50% from the median, two of the results were more than 100% away.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 101, 102, 103, 113, 115, 116, 118 and 131.

Cl-POP 2: Lab. 101, 102, 103, 115, 118, 121 and 131.

α -HCH

Sixteen laboratories took part in the analyses and thirteen participants had two results above the detection limit. The concentrations were similar to those for γ-HCH, and the results both above and lower than the detection limits were quite similar to those of γ-HCH. Laboratory 108 correctly found one of the two concentrations lower than the detection limit while the second was reported far too high, for α-HCH as well as for γ-HCH. Seven laboratories reported one or more results more than 50% from the median, and two results were outside 100% from the median. Laboratories 118, 131, and 115 had both results less than 10% from the median, while laboratories 101 and 103 had results less than 20% from the median.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 101, 103, 115, 118 and 131.

Cl-POP 2: Lab. 101, 102, 103, 109, 115, 118 and 131.

HCB

Sixteen laboratories returned HCB results and one of the participants could not detect this compound. Four participants obtained both results not more than 10% from the median; laboratories 118, 103, 116, and 105. The first two of these had obtained results in the first part that were less than five per cent from the expected. Two laboratories were more than 100% away from the median, while these laboratories both had results less than five per cent from the expected in the first part of the comparison. Another four laboratories had one or more results at least 50% from the median.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 101, 102, 103, 105, 113, 116, 118 and 131.

Cl-POP 2: Lab. 101, 102, 103, 104, 105, 113, 116 and 118.

PCB

Sixteen laboratories took part in the analyses of the seven PCBs; PCB-28, -52, -101, -118, -138, -153, and -180. Laboratory 123 returned, however, results from PCB-28, -52, -101 and -138 only.

The median concentrations of the first three compounds were in the range 3-8 pg/m3 (Table 8). PCB-180 had concentrations at about 1 pg/m3 while the remaining three compounds had medians about 1-2 pg/m3. The concentrations could be typical for northern Europe, but could be in the lower end for the central continental parts (Table 10).

Three laboratories reported all seven compounds lower than their detection limits; 102, 110, and 116. Laboratory 108 reported the concentration to be below the detection limit in all but one case where PCB-153 was reported at a far too high concentration. This was also the case for α- and γ-HCH in the same sample. Laboratory 113 found interfering substances in the measurements of PCB-28, and -52, and laboratory 101 could not detect PCB-101, -118, -138, -153, and PCB-180.

The results for the PCBs given below and in the Annex 5 in general reveal that two laboratories were able to obtain both results better than 10% for more than one PCB; laboratories 103 and 118. Laboratories 118, 131, and 105 all had both results for five of the seven PCBs better than 20% from the median. Laboratories 115 and 103 had four of the seven PCBs within 20%.

In the other end of the scale one laboratory had at least one result more than 100% from the median for all seven PCBs, and a different laboratory had six results at more than 100% from the median. Six laboratories had at least one result more than 100% from the median, and three other laboratories had at least one result deviating 50% or more from the median.

Laboratory 105 commented that use of insufficiently cleaned silica may have given PCB-results up to 20% too high.

PCB-28 (2,4,4'-TriCB)

Three laboratories obtained very good results in this second part; laboratories 118, 101 and103. The results from laboratories 115 and 104 were likewise near the median. Two laboratories reported concentrations more than 100% from the median, both had however obtained very good results in the first part of the comparison. Two other participants reported both results more than 50% higher than the median.

Laboratories inside  20% of the median:

Cl-POP 1 and Cl-POP 2: Lab. 101, 103, 104, 115, 118 and 131.

PCB-52 (2,2',5,5'-TetraCB)

Laboratories 103, 104, 131, and 115 had results very near the median. Two of the sixteen laboratories had both results more than 100% from the median, and one laboratory had both results more than 50% from the median.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 101, 103, 104, 115 and 131.

Cl-POP 2: Lab. 103, 104, 115 and 131.

PCB-101 (2,2',4,5,5'-PentaCB)

Laboratories 113, 103, 118 and 131 had results very close to the median. Three laboratories obtained results more than 100% from the median, while three other laboratories had results at about 50% from the median. Laboratories 109 and 104 that had around 50% deviation from the median, obtained results at about 10% from the expected in the first part of the comparison.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 103, 109, 113, 115, 118 and 131.

Cl-POP 2: Lab. 103, 113, 115, 118, 129A and 131.

PCB-118 (2,3',4,4',5-PentaCB)

The laboratories with results near the median were 104, 109, 115, 118 and 131. Three laboratories had results more than 100% from the median while two other laboratories obtained results that were more than 50% from the median.

Laboratories inside  20% of the median:

Cl-POP 1 and Cl-POP 2: Lab. 104, 109, 115, 118 and 131.

PCB-138 (2,2',3,4,4',5'-HexaCB)

Three laboratories had both results less than 10% from the median; 104, 118 and 131. Three laboratories had results more than 100% from the median, and three others had results outside 50%.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 103, 104, 109, 118, 129A and 131.

Cl-POP 2: Lab. 103, 104, 118, and 131.

PCB-153 (2,2',4,4',5,5'-HexaCB)

Seven single results were less than 10% from the median. Four laboratories had one or two results more than 100% from the median.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 103, 104, 109, 118, 129A and 131.

Cl-POP 2: Lab. 104, 118 and 131.

PCB-180 (2,2',3,4,4',5,5'-HeptaCB)

No laboratories had results less than 20 per cent from the median for both samples. Two of the laboratories that obtained the best result for the other PCBs were 26% at most from the median. As seen from Table 8 above ten laboratories only were able to obtain results above their detection limits. Three laboratories deviated more than 100% from the medians while another laboratory had results more than 50% from the median.

Laboratories inside  20% of the median:

Cl-POP 1: Lab. 103, 104 and 118.

Cl-POP 2: None.

## Results for PAH round 2

The results from the second part of the comparison were evaluated by comparison with the medians of the laboratory results. Table 11 gives the medians and number of laboratories that found results above their detection limit.

Table 12 shows concentration values for the PAHs in the sample in round 1 and in the two samples in round 2. The values for round 2 are calculated from the median for each compound, assuming a 500 m3 sample volume and a final sample extract volume of 100l after clean-up and pre-concentration.

Table 13 give a few simple statistics on the PAH concentrations measured at two sites in Europe in year 2000 (Berg et al., 2002).

Figure 3 and Figure 4 give the naphthalene results from extract 1. Annex 6 presents the results for all laboratories and compounds in similar figures giving the deviation from the median concentration in concentration units, i.e. ng/m3, and as per cent deviation from the median.

Table 11:	Medians and number of laboratories giving results above their detection limits. Medians in ng/m3.

| Compound | 1st sample |  | 2nd sample |  |
| --- | --- | --- | --- | --- |
|  | Median | Number of laboratories | Median | Number of laboratories |
| Naphthalene | 9.63 | 11 | 0.780 | 12 |
| Acenaphthylene | 3.92 | 11 | 0.281 | 12 |
| Acenaphthene | 1.13 | 12 | 0.315 | 13 |
| Fluorene | 4.31 | 12 | 2.61 | 13 |
| Phenanthrene | 8.23 | 11 | 6.68 | 13 |
| Anthracene | 0.800 | 12 | 0.48 | 13 |
| Fluoranthene | 2.75 | 12 | 2.25 | 13 |
| Pyrene | 2.35 | 12 | 1.87 | 13 |
| Benzo[a]anthracene | 0.420 | 12 | 0.270 | 12 |
| Chrysene | 0.780 | 13 | 0.536 | 13 |
| Benzo[b]fluoranthene | 0.507 | 11 | 0.389 | 12 |
| Benzo[k]fluoranthene | 0.226 | 11 | 0.177 | 12 |
| Benzo[a]pyrene | 0.365 | 13 | 0.240 | 14 |
| Indeno[1,2,3-cd]pyrene | 0.402 | 13 | 0.305 | 13 |
| Dibenzo[a,h]anthracene | 0.050 | 10 | 0.042 | 12 |
| Benzo[ghi]perylene | 0.460 | 13 | 0.355 | 14 |
| Biphenyl | 4.62 | 6 | 1.16 | 5 |
| Perylene | 0.0345 | 5 | 0.0200 | 5 |
| Benzo[e]pyrene | 0.409 | 8 | 0.317 | 8 |

Table 12:	Comparison of PAH concentrations in the standard analysed in round 1 and in the samples of round 2.

| Compound | Round 1 Theoretical ng/l | Round 2 PAH1 median ng/l | Round 2 PAH2 median ng/l |
| --- | --- | --- | --- |
| Naphtalene | 8.3 | 48.2 | 3.9 |
| Acenaphthylene | 2.0 | 19.6 | 1.4 |
| Acenaphthene | 4.2 | 5.7 | 1.6 |
| Fluorene | 4.0 | 21.6 | 13.0 |
| Phenanthrene | 8.3 | 41.1 | 33.4 |
| Anthracene | 0.4 | 4.0 | 2.4 |
| Fluoranthene | 4.2 | 13.8 | 11.2 |
| Pyrene | 3.3 | 11.8 | 9.4 |
| Benzo[a]anthracene | 0.6 | 2.1 | 1.4 |
| Chrysene | 1.7 | 3.9 | 2.7 |
| Benzo[b]fluoranthene | 2.2 | 2.5 | 1.9 |
| Benzo[k]fluoranthene | 0.5 | 1.1 | 0.9 |
| Benzo[a]pyrene | 0.4 | 1.8 | 1.2 |
| Indeno[1,2,3-cd]pyrene | 1.3 | 2.0 | 1.5 |
| Dibenzo[a,h]anthracene | 0.2 | 0.3 | 0.2 |
| Benzo[ghi]perylene | 1.0 | 2.3 | 1.8 |
| Biphenyl | 8.0 | 23.1 | 5.8 |
| Perylene | 1.2 | 0.2 | 0.1 |
| Benzo[e]pyrene | 0.4 | 2.0 | 1.6 |

Table 13:	Air concentrations of PAH at sites in the Czech Republic and Sweden in 2000. Results from weekly and monthly samples. PAH unit ng/m3.

| PAH | Košetice CZ0003R |  |  | Rörvik SE0002R |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Min | 50% | Max | Min | 50% | Max |
|  | Weekly samples |  |  | Monthly samples |  |  |
|  | ng/m3 |  |  | ng/m3 |  |  |
| Acenaphtene | 0.01 | 0.08 | 1.51 |  |  |  |
| Phenanthrene | 0.71 | 3.31 | 18.56 | 0.27 | 1.22 | 3.05 |
| Anthracene | 0.01 | 0.06 | 0.69 | 0.0 | 0.02 | 0.07 |
| Pyrene | 0.10 | 0.51 | 5.76 | 0.04 | 0.29 | 0.78 |
| Benzo[a]anthracene | 0.01 | 0.06 | 2.23 | 0.01 | 0.11 | 0.24 |
| Benzo[a]pyrene | 0.005 | 0.059 | 1.757 | 0.005 | 0.059 | 0.267 |
| Indeno[1,2,3-cd]pyrene | 0.01 | 0.07 | 2.06 | 0.04 | 0.07 | 0.37 |
| Benzo[ghi]perylene |  |  |  | 0.01 | 0.05 | 0.32 |

Figure 3:	Results of all naphthalene measurements in extract 1 from the second part of the comparison in ng/m3.

Figure 4:	Deviation in percent from the median of all naphthalene measurements in extract 1 of the second part of the comparison.

Comments to the individual compounds are given below.

Naphthalene

The concentrations in the two extracts differed by a factor of 10, the median concentrations were about 9.6 and 0.8 ng/m3. Twelve participants took part in the measurements of naphthalene and three laboratories obtained both results within 20% from the median; laboratories 105, 116,131. Laboratory 105's results were less than 10% away from the median, and in the first part of the comparison this laboratory was less than 5% from the expected. The two laboratories 116 and 131 did not take part in the first part of the exercise. Three laboratories missed the medians with more than 50%, which may seem surprising since their performance in the first part was very good. One laboratory gave both results more than 100% from the medians.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 105, 116 and 131.

PAH 2: Lab. 105, 110, 116, 129A and 131.

Acenaphthylene

The results of the acenaphthylene measurements in the first part of the comparison were among the least satisfactory in the exercise. The ratio between the acenaphthylene concentrations in the two extracts in this part of the exercise was similar to that of naphthalene, but the concentrations were lower; about 3.9 and 0.3 ng/m3. Three laboratories of the twelve participants obtained both results less than 10% from the median; laboratories 129A, 116, and 101. Neither 129A nor 101 obtained good results in the first part, results they shared with the three laboratories with results more than 50% from the median. Two laboratories gave respectively one and two results more than 100% from the medians.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 109, 116, 123 and 129A.

PAH 2: Lab. 101, 102, 103, 116, 129A and 131.

Acenaphthene

The median concentrations of acenaphthene were about 1.1 and 0.3 ng/m3. These were rather typical weekly concentrations from the central part of Europe (Košetice in the Czech Rep., Table 13) in 2000. Thirteen laboratories measured this compound. Two laboratories, 116 and 131 had both results closer than 10% from the median. Three laboratories had five results outside 50%, and four of these measurements were more than 200% from the median. One of the laboratories outside 200% had excellent results in the first part of the exercise.

Laboratories inside  20% of the median:

PAH 1: Lab. 107, 116 and 131.

PAH 2: Lab. 116, 129A and 131.

Fluorene

Thirteen participants have analysed both extracts. Laboratory 102 measured PAH in the low concentration extract only. The median concentrations of fluorene were approximately 4.3 and 2.6 ng/m3. Three participants had both results less than 10% from the medians; laboratories 103, 109, and 116. Lab. 109 had excellent results in the first part as well while Lab. 103 was about 24% from the expected. Three laboratories had results more than 50% away, and one had very large deviations, but very good results in the first part. This is in line with the results above; a good ability to calibrate and analyse a standard is unfortunately no guarantee for good results on real extracts.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 109, 116 and 131.

PAH 2: Lab. 103, 107, 109 and 116.

Phenanthrene

The participants in this measurement were as for fluorene. The phenanthrene medians reached about 6.7 and 8.2 ng/m3. This seems to be higher than one would expect to find in northern Europe but could be rather typical for the central parts (Table 13). The results were without outliers as for some of the compounds above, one laboratory got results between 50% and 60%, and these were the most deviating results. Two laboratories were within 10% from the median; 109 and 131. Laboratory 109 had excellent results in the first part. Another three laboratories found phenanthrene in both extracts less than 20% away from the medians, these were laboratories 103, 105 HPLC and 107 that had similar good results in part 1.

Laboratories inside  20% of the median:

PAH 1: Lab. 103, 105, 107, 109, 129A and 131.

PAH 2: Lab. 102, 103, 105, 107, 108, 109 and 131.

Anthracene

The anthracene concentrations were both low, about 0.5 and 0.8 ng/m3. This is higher than one would expect at regional sites in the northern part of Europe and probably in the high end of the Continental regional concentrations (Table 13). This may be the reason for some of the apparent large deviations from the median. It is, however, important to bear in mind that the small concentrations easily give high deviations in per cent for fairly small concentrations differences. Fourteen laboratories measured both concentrations. Two laboratories found concentrations within 10%; laboratories 131 and 110. Three laboratories had outliers more than 300% from the median, and one had results 60–80% too high.

Laboratories inside  20% of the median:

PAH 1: Lab. 103, 107, 110 and 131.

PAH 2: Lab. 102, 103, 110 and 131.

Fluoranthene

Thirteen participants returned results for both extracts and one participant had a too high detection limit to measure these concentrations. The median concentrations were higher than anthracene’s concentrations, 2.8 and 2.3 ng/m3. Five laboratories obtained both results within 10 % from the median and two more within 20%. These were laboratories 131, 109, 129A, 101, and 103 and 116. No results were outside 100% from the median but two laboratories had results larger than 50%.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 107, 109, 116, 129A and 131.

PAH 2: Lab. 101, 102, 103, 116, 129A and 131.

Pyrene

Pyrene was one of the compounds with the largest number of very good results in the first part of the comparison, and the results from the second part were also good. Fourteen laboratories carried out measurements of both extracts; laboratory 108 had, however, a too high detection limit as for fluoranthene. The medians were 2.4 and 1.9 ng/m3, both concentrations close to this detection limit. Two laboratories gave results within 10% from the median; laboratories 113 and 131. Three results from two laboratories were between 50% and 100% from the median.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 107, 113, 129A and 131.

PAH 2: Lab. 101, 103, 110, 113, 116, 129A and 131.

Benzo[a]anthracene

The concentrations of benzo[a]anthracene were rather low, about 0.4 and 0.3 ng/m3, and could be rather typical for central parts of Europe, but higher than expected in the north (e.g. Table 13). The results in the second part of the comparison were nevertheless good. Twelve participants analysed both extracts and five laboratories returned both results better than 20%, one of which were better than 10% for both samples; laboratory 101. Four single results were more than 50% from the median, two of which were larger than 200%.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 104, 110, 113, 123, 129A and 131.

PAH 2: Lab. 101, 103, 104, 116, 129A and 131.

Chrysene

The concentrations were low, about 0.8 and 0.5 ng/m3. Thirteen participants analysed both samples and five laboratories obtained both results better than 20%; laboratories 103, 113, 131, 129A, and 101. The two first laboratories were not more than 10% away. Four results from two laboratories were more than 100% from the median. GC-methods may produce high results compared to HPLC-methods because of co-elution with triphenylene.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 113, 129A and 131.

PAH 2: Lab. 101, 102, 103, 113, 116, 129A and 131.

Benzo[b]fluoranthene

The laboratories obtained very good results for this compound in round one. The concentrations in round two corresponds to about 0.5 and 0.4 ng/m3, and eleven participants analysed both samples. Three laboratories reported both results closer than 10% from the median while another participant was 11% and 0% away in the two extracts. The laboratories were laboratories 116, 129A, 103 and 101. One participant was about 90% off in both two results while one laboratory was 60% and 80% away. GC-methods may produce high results compared to HPLC-methods because of co-elution with benzo[j]fluoranthene and benzo[k]fluoranthene. Laboratory 131 reported the sum of benzo[b]fluoranthene and benzo[k]fluoranthene.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 110, 116 and 129A.

PAH 2: Lab. 101, 103, 105, 116 and 129A.

Benzo[k]fluoranthene

Eleven participants analysed both extracts that had medium concentrations about 0.2 ng/m3. Two participants had both results closer than 10%, and two other laboratories measured concentrations not more than 20% off; laboratories 104 and 101, and 103 and 110. Two laboratories measured far too high concentrations. The results for this component and the previous one were not better than those for chrysene and benzo[a]anthracene in contrast to the results from the first part of this exercise where benzo[b]fluoranthene and benzo[k]fluoranthene obtained the largest number of the very best results. GC-methods may produce high results compared to HPLC-methods because of co-elution with benzo[b]fluoranthene and benzo[j]fluoranthene

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 104 and 110.

PAH 2: Lab. 101, 102, 103, 104, 105 and 110.

Benzo[a]pyrene

Thirteen laboratories returned results from both extracts. The concentrations of benzo[a]pyrene were about 0.4 and 0.2 ng/m3. Only laboratory 101 obtained both results less than 10% from the median. One of the two laboratories with strongly deviating results in the analysis of the previous compound had far too high results for this compound as well.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 110, 113, 129A and 131.

PAH 2: Lab. 101, 102, 103, 110, 116, 129A and 131.

Indeno[1,2,3-cd]pyrene

Twelve laboratories analysed both extracts, and three laboratories obtained results less than 10% from the median; laboratories  131, 110 and 129A. Three results from two laboratories were far too high and one results was about 50% too high. The concentrations indeno[1,2,3-cd]pyrene were about 0.3 and 0.4 ng/m3. In Central Europe much higher concentrations may be found (Table 13), but the concentrations are generally lower than this. The concentrations in northern parts of Europe are generally much lower.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 105, 109, 110, 113, 129A and 131.

PAH 2: Lab. 101, 102, 105, 109, 110, 113, 129A and 131.

Dibenzo[a,h]anthracene

Thirteen laboratories returned results for this compound. The median concentrations were quite low, about 0.05 ng/m3, but the results were good with little scatter except from four large outliers. Four laboratories returned both results within 10% from the median; laboratories 101, 131, 105 and 103. In addition to the four large outliers three single results were somewhat higher than 50%. One of the participants with a large deviating result had a too large concentration also in the first part while the other three laboratories obtained results within 20% from the expected. GC-methods may produce high results compared to HPLC-methods because of co-elution with dibenzo[a,c]anthracene.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 105, 113 and 131.

PAH 2: Lab. 101, 103, 105, 116, 129A and 131.

Benzo[ghi]perylene

The medians obtained were between 0.3 and 0.5 ng/m3. Thirteen participants analysed both samples. The concentrations were similar to the highest monthly concentration measured at Rörvik (Table 13) in year 2000. Four participants had both results within 10%; laboratories 131, 113, 101 and 103. Three laboratories had both results more than 50% from the median, all but one single result was <100% from the median.

Laboratories inside  20% of the median:

PAH 1: Lab. 101, 103, 105, 107, 113, 129A and 131.

PAH 2: Lab. 101, 103, 105, 113, 116, 129A and 131.

Biphenyl and perylene

Only five and six laboratories analysed both samples for biphenyl and perylene respectively. There were three single results between 50% and 100% from the median found for biphenyl, the remaining concentrations being less than 50%. This is in contrast to perylene results where it is seen that one laboratory has two very large outliers, and one laboratory had one large deviating result. The concentrations for perylene were, however, very low 0.02-0.04 ng/m3 only. The biphenyl concentrations were higher, 1-5 ng/m3.

Benzo[e]pyrene

The benzo[e]pyrene median concentrations were about 0.3 to 0.4 ng/m3 that were measured by eight participants only. The results for benzo[e]pyrene are consistent in both samples and have systematic deviations from the median. Results reported by one laboratory were larger than 100% from the median, while those from two other laboratories were between 50% and 100%. The remaining results were mostly below thirty per cent, and results from three laboratories were better than 20%.

Laboratories inside  20% of the median:

PAH 1 and PAH 2: Lab. 101, 129A and 131.

# Conclusion

An international laboratory comparison was conducted in order to obtain information on the comparability of results from monitoring data of POPs in air.

The comparison was a two-step exercise:

Round 1: Analysis of a mixture of standards of known composition but with unknown concentrations.

Round 2: Analysis of two raw-extracts from filter and gas phase adsorbents after high volume air sampling.

Two groups of POPs were investigated during the exercise: Organochlorine compounds including pesticides, hexachlorobenzene (HCB) and polychlorinated biphenyls (PCBs) and polycyclic aromatic hydrocarbons (PAHs).

The organochlorine compounds analysed were:

p,p’-DDT, p,p’-DDE, -Chlordane, -Chlordane, Hexachlorocyclohexane (HCH): γ-HCH and α-HCH, HCB, PCB 28, PCB 52, PCB 101, PCB 118, PCB 138, PCB 153 and PCB 180.

The PAHs analysed were:

Naphthalene, Acenaphthylene, Acenaphthene, Fluorene, Phenanthrene, Anthracene, Fluoranthene, Pyrene, Benzo[a]anthracene, Chrysene, Benzo[b]fluoranthene, Benzo[k]fluoranthene, Benzo[a]pyrene, Indeno[1,2,3-cd]pyrene, Dibenzo[a,h]anthracene, Benzo[ghi]perylene, Biphenyl, Perylene and Benzo[e]pyrene.

Some laboratories analysed only one compound group.

Results from round 1 are reported as: concentration found in the sample (as pg/l or ng/l) and as deviation from expected value (in %).

The total number of laboratories involved was 21 of which 16 took part in round 1. The results from the first part show that all but one laboratory had at least one result not more than 20 % from the expected, for PAH and for the organochlorine compounds.

Seven laboratories had all, but 3–4 PAH results, better than 20% from the expected, and five laboratories had all except one organochlorine results better than 20%.

Six laboratories had no PAH results, and five laboratories no organochlorine results, more than 50% from the expected in the first part of the comparison.

Results from round 2 are reported as sample concentration found (as pg/m3 or ng/m3, assuming a sample volume of 500 m3) and as deviation from median (in %).

Eighteen laboratories took part in round 2. Four laboratories had more than twenty single results of PAH within 20% from the median, and two laboratories obtained a similar result for the organochlorine compounds.

The second round showed the importance of using an analytical method and instrumentation able to cope with possible interferences without loss of analyte even while working at low concentrations, while round 1 mainly was a calibration exercise.

In round 2, organochlorine compounds, three laboratories accounted for about 66% of the results deviating from the median with 50% or more. Three laboratories accounted for 53% of the PAH results deviating from the median with 50% or more.

The results give an indication of the laboratory accuracy that can be expected for POPs in air. In an actual monitoring situation the accuracy would also be influenced by sampling method (sampling time, flow, filter materials, sampler type), as well as the analytical precision.

# References

Berg, T., Hjellbrekke, A.G. and Larsen, R. (2002) Heavy metals and POPs within the EMEP region 2000. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 9/2002).

Berg, T., Hjellbrekke, A.G. and Skjelmoen, J.E. (1996) Heavy metals and POPs within the ECE region. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 8/96).

Lükewille, A. (ed.) (1998) EMEP expert meeting on measurements of persistent organic pollutants (POP) in air and precipitation. Kjeller, Norwegian Institute for Air Research (EMEP/CCC-Report 8/98).

- List of participants

| Lab no. | Country | Laboratories | Contact(s) |  |
| --- | --- | --- | --- | --- |
| 101 | AUSTRIA | Federal Environment Agency – Austria Analytics Spittelauer Lände 5 AT-1090 VIENNA Phone: +43 1 3130 45216 Fax: +43 1 31 304 5222 | Gundi Lorbeer lorbeer@ubavie.gv.at | Werner Hartl hartl@ubavie.at |
| 102 | BELGIUM | Vlaamse Milieumaatschappij Krijgslaan 281, blok S 2 BE-9000 GENT Phone: +32 9 264 4404 Fax: +32 9 264 4981 | Eric De Wulf e.dewulf@vmm.be | Eric Wauters e.wauters@vmm.be |
| 103 | CANADA | Meteorological Service of Canada Organics Analysis Laboratory 4905 Dufferin Street Downsview ONTARIO M3H 5T4 Phone: +1 416 739 4601 Fax: +1 416 739 5708 | Ken Brice ken.brice@ec.gc.ca |  |
| 104 | CANADA | Freshwater Institute Department of Fisheries and Oceans 501 University Crescent WINNIPEG MB, R3T 2N6 Phone: +1 204 984 6761 Fax: +1 204 984 2403 | Gary Stern sterng@dfo-mpo.gc.ca |  |
| 105 | CZECH REPUBLIC | RECETOX – TOCOEN & Associates Kamenice 126/3 CZ-625 00 BRNO Phone: +420 5 4712 1401 Fax: +420 5 4712 1431 | Ivan Holoubek holoubek@recetox.muni.cz | Jana Klánová klanova@recetox.muni.cz |
| 107 | FRANCE | Laboratoire National d’Essais CMI 37 1, rue Gaston Boisser FR-75724 PARIS Phone: +33 1 4043 3853 Fax: +33 1 4043 3737 | Béatrice Lalere beatrice.lalere@lne.fr |  |
| 108 | FRANCE | Laboratoires Wolff-Environment Waters, Air, Soils and Solid Wastes Analysis Z.I. Saint-Guénault 7, rue Jean Mermoz FR-91031 EVRY Phone: +33 1 6936 5189 Fax: +33 1 6936 5188 | Claude Hennequin claude_hennequin@sgs.com | Emmanuelle Drab emanuelle_drab@sgs.com |
| 109 | FRANCE | Laboratoire CALYDRA 13, rue Férréol Prézelin BP39 FR-44560 PAIMBOEUF Phone: Fax: +33 2 4064 0401 | Eric Bernardy lab@calydra.fr |  |
| 110 | FRANCE | Laboratoire LARA 75, voie Du Toec FR-33300 TOULOUSE Phone: Fax: +33 5 6161 0931 | Stéphanie Alonso s.alonso@lara-laboratoire.com |  |
| 113 | GERMANY | Umweltbundesamt II 3.6 – Labor für Wasseranalytik Bismarckplatz 1 DE-14193 BERLIN Phone: +49 30 8903 2665 Fax: +49 30 8903 2965 | Manfred Harnisch manfred.harnisch@uba.de | George Sawal george.sawal@uba.de |
| 115 | GERMANY | Alfred-Wegener-Institut Postfach 120161 DE-27515 BREMERHAVEN Phone: Fax: | Sönke Lakaschus slakaschus@awi-bremerhaven.de |  |
| 116 | GERMANY | IfE Leipzig Torgauer Str. 116 DE-04347 LEIPZIG Phone: Fax: | Knut Niebergall hanrieder@ife-le.de |  |
| 118 | ICELAND | University of Iceland Department of Pharmacology P.O. Box 8216 IS-128 REYKJAVIK Phone: +354 5680 866 Fax: +354 5680 872 | Kristin Olafsdottir stinaola@hi.is |  |
| 120 | IRELAND | Environmental Protection Agency Pottery Road Dun Laoghaire CO. DUBLIN Phone: +353 1 2852 122 Fax: +353 1 2851 766 | Ciaran O'Donnell c.odonnell@epa.ie |  |
| 121 | LITHUANIA | Joint Research Centre Environmental Ministry of the Republic of Lithuania A. Juozapavičiaus 9 LT-2600 VILNIUS Phone: +370 2 723 202 Fax: +370 2 723 202 | Nijole Striupkuviene nijole.striupkuviene@nt.gamta.lt |  |
| 122 | LITHUANIA | Institute of Physics A. Gostauto 12 LT-2600 VILNIUS Phone: +370 2 651 854 Fax: +370 2 235 182 / 617 070 | Audronė – Regina Milukaitė amk@ktl.mii.lt |  |
| 123 | THE FORMER YUGOSLAV REPUBLIC OF MACEDONIA | Republic Hydrometeorological Institute Skubi bb MK-1000 SKOPJE Phone: +389 91 397 103 Fax: +389 91 397 118 | Mirko Cvetkovski | Radmila Simeva radmila.simeva@yahoo.com |
| 129 | UNITED KINGDOM | AEA Environmental National Technology Centre Culham ABINGDON OX14 3DB Phone: +44 1235 463 084 Fax: +44 1235 463 038 | Peter Coleman peter.coleman@aeat.co.uk |  |
| 130 | YUGOSLAVIA | Federal Hydrometeorological Institute Bircaninova 6 P.O. Box 604 YU-11001 BELGRADE Phone: +381 11 646 555 Fax: +381 11 646 369 | Tanja Vukovic tvukovic@meteo.yu |  |
| 131 | NORWAY | Norwegian Institute for Air Research P.O. Box 100 NO-2027 KJELLER Phone: +47 6389 8000 Fax: +47 6389 8050 | Jan Schaug jan.schaug@nilu.no | Stein Manø stein.mano@nilu.no |

- Methods

Overview of instrumentation used

PAH:

| Lab no. | GC/FID | GC/MS | HPLC | TLC |
| --- | --- | --- | --- | --- |
| 101 |  | X |  |  |
| 102 |  |  | X |  |
| 103 |  |  | X |  |
| 104 |  | X |  |  |
| 105 |  | X | X |  |
| 107 |  | X |  |  |
| 108 |  |  | X |  |
| 109 | X |  | X |  |
| 110 |  |  | X |  |
| 113 |  |  | X |  |
| 116 |  | X | X |  |
| 122 |  |  |  | X |
| 123 | X |  |  |  |
| 129A |  | X |  |  |
| 129L |  | X |  | X |
| 130 |  | X |  |  |
| 131 |  | X |  |  |

Chlorinated compounds:

| Lab no. | GC/ECD | GC/MS |
| --- | --- | --- |
| 101 |  | X |
| 102 | X |  |
| 103 |  | X |
| 104 | X |  |
| 105 | X |  |
| 108 | X |  |
| 109 | X |  |
| 110 |  | X |
| 113 |  | X |
| 115 |  | X |
| 116 | X |  |
| 118 | X |  |
| 120 |  | X |
| 121 | X |  |
| 123 | X |  |
| 129A | X |  |
| 129L |  | X |
| 131 |  | X |

Lab 101

| Method description PAH |  | Method description Cl - POP |  |
| --- | --- | --- | --- |
| Gas-chromatography |  | Gas-chromatography |  |
| Column phase | DB 5 MS | Column phase | DB 5 MS |
| Length (m) | 60 m | Length (m) | 60 m |
| Internal diameter (mm) | 0,25 | Internal diameter (mm) | 0,25 |
| Film thickness (µm) | 0,25 | Film thickness (µm) | 0,25 |
| Type of injector | Splitless | Type of injector | Splitless |
| Injection volume (µl) | 2 | Injection volume (µl) | 2 |
| Injection temperature (°C) | 270 | Injection temperature (°C) | 270 |
| Detector |  | Detector |  |
| Detector type | Agilent MSD | Detector type | Thermoquest Voyager MS |
| Detector temperature (°C) | 280 °C | Detector temperature (°C) | 260 |

Lab 102

PAHs: HPLC with programmed fluorescence detection (UV for Acenaftylen). Solvent replaced by acetonitrile.

Chlorinated compounds: GC/ECD on HT-8 and CP-SIL 19 columns.

Lab 103

PAHs:

Each sample was diluted to 2 mL in i-octane, using half of that volume in the analysis.

Cleanup was by silica SPE using Waters cartridges on a Hamilton Microlab Workstation, with elution by dichloromethane and solvent exchange into 1mL of acetonitrile.

Instrumental analysis was by reverse-phase gradient HPLC-FLD using external standardisation. Multiple dilutions of up to 50X were necessary to bring the analyte levels into the usual calibration range of the method used for IADN. The optimum dilution range required was PAH dependent. The relevant vials were injected in triplicate to generate a measure of instrumental precision.

Detection limits were calculated from the estimated instrumental detection limit and converted to "atmospheric equivalent detection limits" using assumptions of the same conditions for the samples in question i.e sample split of 2, final vial volume of 1 mL, air volume nominally 500 m3.

Perylene is not uniquely determined in the MSC method, but (based on the results from EMEP Round 1 and other retention data) it is believed that this species probably co-elutes with benzo[b]fluoranthene and will produce a positive bias in the values reported if it is present in the ambient sample extracts provided.

Chlorinated compounds:

Each of the ampoules was accurately diluted to 2 mL in i-octane, with half of that volume being subjected to analysis.

Florisil column chromatography (60-100 mesh, calcined at 600oC and then deactivated with approximately 2% w/w water) was used to remove interferences and fractionate the sample into three eluates: F1 (eluted with hexane) contained most PCBs plus HCB, heptachlor, aldrin, p,p'-DDE, o,p'-DDT, mirex, photomirex along with t-nonachlor (partial) and p,p'-DDT (partial) and α-chlordane (minor contribution); F2 (eluted with 15% DCM/85% hexane) contained coplanar PCBs plus α-HCH, β-HCH, γ-HCH, oxychlordane, α-chlordane, γ-chlordane, o,p'-DDD, p,p'-DDD along with t-nonachlor (partial) and p,p'-DDT (partial); F3 (eluted with 60% DCM/40% hexane and then DCM alone) contained dieldrin, endrin, heptachlor epoxide, methoxychlor, α-endosulfan, β-endosulfan and δ-HCH.  The final extract fractions were concentrated by Turbovap nitrogen blowdown, accurately adjusted to a volume of 1 cm3 and then transferred to autosampler vials fitted with PTFE septa to eliminate silicone interferences.  Analysis was performed using a HP5890 GC equipped with dual 63Ni ECDs and dual heated splitless/split injection ports. A 60 m x 0.25 mm i.d. (df = 0.25 μm) DB5 column (J&W Scientific) was used for primary analysis (800C for 2 min, 150C/min to 1600C then 2.50C/min to 2650C and hold for 20 min). A 30 m x 0.25 mm i.d. (df = 0.25 μm) DB-17 column operated under the same temperature program was used to provide confirmatory analysis and/or enhanced separation for certain of the co-elutions on the DB5 column. Splitless-mode injections of 1μL (purge on at 2.0min) were made separately onto both columns using an autosampler/autoinjector system (HP 7673 series). Ultra high purity (UHP) He (Praxair) was used as carrier gas in constant-flow mode, with initial column flow-rates of approximately 1.3 mL/min.  Injection port and detector were held at 2000C and 3500C, respectively. ECD make-up gas was UHP N2 (Praxair) at 70 mL/min. External standard calibrated methods were used for quantitation. Note that in the analytical sequence, frequent "bracketing" of the samples with injections of standards was employed to track and correct for response drift.

Lab 104

PAHs: Sample injected on Hewlett Packard 6890 GC with 5973 MSD. Column: J&W DB-5MS 30 x 0.25m.

Chlorinated compounds: The samples were cleaned up and fractionated on a 8g Florisil column (wet packed in hexane). The hexane fraction included PCBs, most HCB and pp-DDE while the 15/85 hexane/DCM fraction included HCHs, chlordanes and pp-DDT. Samples were reduced to 0.5 mL for injection on a GC equipped with an ECD. Detection limits are roughly 0.1 pg/m3 (for PCB 28) or lower.

Lab 105

First Round of the EMEP POP Laboratory Comparison

Sample:	PAH U No. 47 (NILU, Norway)

Remark:

Lifetime of multiplier is over and, therefore, not all tuning parameters are fulfilled. Detector sensitivity is lower.

GC-MS Equipment:	GC Model HP 6890 with Mass Selective Detector HP 5972A, Autosampler HP 6890, HP Enhanced ChemStation G1701AA Version A.03.00 (Hewlett-Packard, FRG).

GC Column:	DB-5ms, 58.5 m x 0.25 mm x 0.25 µm film thickness (J&W, USA)

Carrier Gas:	Helium, purity 99.9995 %, constant flow 1.5 ml/min

Temperatures:	Injector 280°C

Interface 280°C

Temperature programme:	80°C 1 min // 15°C/min to 180°C // 5°C/min to 310°C // 310°C 10 min

Injected Volume:	1 μl splitless

Sample preparation:	200 µl aliquot of PAH U No. 47 solution, 50 µl of internal standard (p-Terphenyl) and 750 µl toluene was added

Standards:	Supelpreme-HC PAH-Mix (SUPELCO, USA), solution in dichloromethane:benzene (50:50)

Biphenyl Pestanal (Riedel-de Haën, FRG), Benzo[e]pyrene (Dr. Ehrenstorfer, FRG), Perylene sublimed (Sigma Aldrich, FRG)

Biphenyl, Benzo[e]pyrene and Perylene solutions in acetonitrile prepared by weighing and dilution.

Calibration:	Six-level calibration in the range from 0.40 to 10.0 ng/μl

Detection:	Mass Selective Detection in Selected Ion Monitoring Mode

| Selected Ion Monitoring |  |  |  |
| --- | --- | --- | --- |
|  | Target m/z | Qualifier m/z |  |
| Terphenyl I.S. | 230 | 228 | 215 |
| Naphthalene | 128 | 129 | 126 |
| Acenaphthylene | 152 | 153 | 150 |
| Acenaphthene | 154 | 153 | 155 |
| Fluorene | 166 | 167 | 164 |
| Phenanthrene | 178 | 179 | 176 |
| Anthracene | 178 | 179 | 176 |
| Fluoranthene | 202 | 203 | 200 |
| Pyrene | 202 | 203 | 200 |
| Benzo[a]anthracene | 228 | 229 | 226 |
| Chrysene | 228 | 229 | 226 |
| Benzo[b]fluoranthene | 252 | 253 | 250 |
| Benzo[k]fluoranthene | 252 | 253 | 250 |
| Benzo[a]pyrene | 252 | 253 | 250 |
| Indeno[1,2,3-cd]pyrene | 276 | 277 | 274 |
| Dibenzo[a,h]anthracene | 278 | 279 | 276 |
| Benzo[ghi]perylene | 276 | 277 | 274 |
| Biphenyl | 154 | 152 | 155 |
| Perylene | 252 | 253 | 250 |
| Benzo[e]pyrene | 252 | 253 | 250 |

First Round of the EMEP POP Laboratory Comparison

HPLC Determination with UV-VIS Diode-Array Detection:

Remark:

Retention times and, also, peak shapes were affected by the presence of i-octane. Other solvents as e.g., methanol or acetonitrile, would be compatible with HPLC determination.

At standard conditions, fluorescence detection would be used in combination with UV-VIS diode-array detection (for acenaphthylene and chromatographically unresolved peaks of benzo[b]fluoranthene and perylene). Recently, our fluorescence detector (HP 1046A) is out of work - waiting for service.

HPLC Equipment:	Model HP 1050 quaternary pump, programmable autosampler, and diodearray detector, model HP 1100 thermostatted column compartment, HP ChemStation Rev. A.07.01 for LCD 3D Systems (all Hewlett-Packard, FRG).

HPLC Column:	Narrow-bore Supelcosil LC PAH, 250 x 2.1 mm I.D. (Supelco, USA) packed with 5-μm octadecyl phase

Guard column – SecurityGuard System with C18 cartridges (Phenomenex, USA)

Column Temperature:	20±0.5°C

Injected Volume:	3.6 μl

Standards:	PAH-Mix 45 (Dr. Ehrenstorfer, FRG), solution in acetonitrile

Biphenyl Pestanal (Riedel-de Haën, FRG),

Biphenyl solutions in acetonitrile prepared by weighing and dilution.

Calibration:	Six-level calibration in the range from 0.60 to 10.0 ng/μl

Reagents:	Individual polycyclic aromatic hydrocarbons (Supelco, USA)

Acetonitrile, HPLC gradient grade (J. T. Baker, The Netherlands)

Water from Millipore Q185

Gradient elution:	Component A:	acetonitrile/water 50:50

Component B:	acetonitrile

Flow:		0.35 ml/min

| Time (min) | % B |
| --- | --- |
| 0 | 0 |
| 5 | 0 |
| 30 | 100 |
| 47 | 100 |

Detection:	Signals at 225 nm (20 nm bandwidth), 245 (15), 280 (20), 297 (26), and425 nm (30) were used for detection as shown in table with results. Reference wavelength 530 nm (50 nm) Spectra taken from 210 to 450 nm

First Round of the EMEP POP Laboratory Comparison

Sample:	Cl-POP U No. 5 (NILU, Norway)

GC-ECD Equipment:	GC Model HP 5890 Series II with Electron Capture Detector, Autosampler HP 7673, HP GC ChemStation Rev A.03.03 (Hewlett-Packard, FRG).

GC Column:	Quadrex 007-5, 40m x 0.18mm x 0.25µm film thickness (Quadrex, USA)

Carrier Gas:	Helium, purity 99.9995 %, constant flow 0.3 ml/min

Make-up gas:	Nitrogen, purity 99.998 %, flow 50 ml/min

Temperatures:	Injector 280°C

Detector 310°C

Temperature programme:	80°C 1 min // 20°C/min to 160°C // 3°C/min to 300°C // 300°C 5 min

Injected Volume:	1 μl splitless

Sample preparation:	400 µl aliquot of Cl-POP U No. 5 solution, 50 µl of internal standard (PCB 155) and 550 µl cyclohexane was added

Standards:	PCB-Mix I (Dr.Ehrenstorfer, FRG), solution in cyclohexane

PCB 118 (ACU Standards) in isooctane

EPA Method 8270A – Chlorinated pesticides (Absolute Standards, USA)

Hexachlorobenzene (SUPELCO, USA) in acetone

Gama-Chlordane and alfa-Chlordane was calibrated using Cl-POP Standard Nr. 54 (NILU, Norway)

Calibration:	Six-level calibration in the range from 10 to 4000 pg/μl

Second Round of the EMEP POP Laboratory Comparison

PAHs Method Description

Samples:	PAH 1 No. 33, PAH 2 No. 4

Sample clean-up:	column chromatography on activated silica (activation 16 hours, 150°C)

fraction 15 ml hexane discarded, fraction 15 ml dichloromethane was concentrated and internal standard (p-Terphenyl) was added

GC-MS Equipment:	GC Model HP 6890 with Mass Selective Detector HP 5972A, Autosampler HP 6890, HP Enhanced ChemStation G1701AA Version A.03.00 (Hewlett-Packard, FRG).

GC Column:	DB-5ms, 58 m x 0.25 mm x 0.25 µm film thickness (J&W, USA)

Carrier Gas:	Helium, purity 99.9995 %, constant flow 1.5 ml/min

Temperatures:	Injector 280°C

Interface 280°C

Temperature program:	80°C 1 min // 15°C/min to 180°C // 5°C/min to 310°C // 310°C 10 min

Injected Volume:	1 μl splitless

Standards:	Supelpreme-HC PAH-Mix (SUPELCO, USA), solution in dichloromethane:benzene (50:50)

Calibration:	Nine-level calibration in the range from 0.001 to 10.0 ng/μl

Detection:	Mass Selective Detection in Selected Ion Monitoring Mode

| Selected Ion Monitoring |  |  |  |
| --- | --- | --- | --- |
|  | Target m/z | Qualifier m/z |  |
| Terphenyl I.S. | 230 | 228 | 215 |
| Naphthalene | 128 | 129 | 126 |
| Acenaphthylene | 152 | 153 | 150 |
| Acenaphthene | 154 | 153 | 155 |
| Fluorene | 166 | 167 | 164 |
| Phenanthrene | 178 | 179 | 176 |
| Anthracene | 178 | 179 | 176 |
| Fluoranthene | 202 | 203 | 200 |
| Pyrene | 202 | 203 | 200 |
| Benzo[a]anthracene | 228 | 229 | 226 |
| Chrysene | 228 | 229 | 226 |
| Benzo[b]fluoranthene | 252 | 253 | 250 |
| Benzo[k]fluoranthene | 252 | 253 | 250 |
| Benzo[a]pyrene | 252 | 253 | 250 |
| Indeno[1,2,3-cd]pyrene | 276 | 277 | 274 |
| Dibenzo[a,h]anthracene | 278 | 279 | 276 |
| Benzo[ghi]perylene | 276 | 277 | 274 |

Cl-POP Method Description

Samples:	Cl-POP 1 No. 54, Cl-POP 2 No. 8

Sample clean-up:	column chromatography on activated silica (activation 16 hours, 150°C) modified 44w% H2SO4,

collected fraction 25 ml hexane:dichloromethane (95:5) was concentrated and internal standard (PCB 155) was added

GC-ECD Equipment:	GC Model HP 5890 Series II with Electron Capture Detector, Autosampler HP 7673, HP GC ChemStation Rev A.03.03 (Hewlett-Packard, FRG).

GC Column:	Quadrex 007-5, 40m x 0.18mm x 0.25µm film thickness (Quadrex, USA)

Carrier Gas:	Helium, purity 99.9995 %, constant flow 0.3 ml/min

Make-up gas:	Nitrogen, purity 99.998 %, flow 54 ml/min

Temperatures:	Injector 280°C

Detector 310°C

Temperature program:	80°C 1 min // 20°C/min to 160°C // 3°C/min to 300°C // 300°C 5 min

Injected Volume:	1 μl splitless

Standards:	PCB-Mix 3 (Dr.Ehrenstorfer, FRG), solution in isooctane

Pesticide-Mix 13 (Dr.Ehrenstorfer, FRG), solution in cyclohexane

Calibration:	Eight-level calibration in the range from 1 to 4000 pg/μl

Lab 107

PAHs round 1:

Isotopic dilution for each compound / GC-MS

Standard:	SRM 2260

Column:	(5% - Phenyl) – methylpolysiloxane, 60 m, 0.25 mm, 0.25 µm

Program temperature:	50°C –2 min

5°C / min  290 °C

290 °C – 30 min

PAHs round 2:

Standard:	SRM 2260

Column:	(5% - Phenyl) – methylpolysiloxane, 60 m, 0.25 mm, 0.25 µm

Program temperature:	50°C –2 :min

5°C / min  290 °C

290 °C – 30 min

Ampoule 13:

The extract was analysed by Isotopic Dilution Gas chromatography / Mass Spectrometry (DI GC/MS).

Ampoule 13:

The ampoule was purified on silica and alumina column and eluted with dichloromethane in hexane (20%). The extract was analysed by DI GC/MS.

Lab 108

PAHs:

The solution is analysed by high performance liquid chromatography. The identification and the quantification of the PAH is made with fluorescence and UV detector.

Column: Vydac, length 150mm, internal diameter 4.6mm

Gradient elution (H2O/CH3CN) constant flow (1ml/min)

Wavelength programming for the excitation  and emission wavelength

UV detection : only for Acenaphtylene.

Limit of detection : 10µg/l.

Uncertainty: around 10%.

Notice that we have co-elution of Benzo (b) fluoranthène and perylene

Chlorinated compounds:

Determination of Organochloro Pesticides :

The solution is analysed by capillary gas chromatography.

GC conditions:

Separation column: Rtx-CLP column, length 60 m, internal diameter 0.25mm, filmthickness 0.25µm.

Oven temperature programmable

Splitless injection

Carrier gas : hydrogen 1ml/min.

Detector : ECD. T = 300°C.

Limit of detection : 5µg/l.

Uncertainty: around 10%

Determination of Polychlorinated biphenyls (PCB) :

The solution is analysed by capillary gas chromatography.

GC conditions:

Separation column: XTI5 (5% phényl)+HT8 (8%phényl) columns, length 50 m, internal diameter 0.25mm, filmthickness 0.25µm.

Oven temperature programmable

Splitless injection

Carrier gas : hydrogen 1ml/min.

Detector :  electron capture detector (ECD). T = 300°C.

The PCB are quantified using an internal standard added to the extract.

Limit of detection : 5µg/l.

Uncertainty: around 10%

Lab 109

PAHs:

Compound		Detection

Naphtalene		GC/FID

Acenaphthylene		GC/FID

Acenaphthene		GC/FID

Fluorene		GC/FID

Phenanthrene		GC/FID

Anthracene		HPLC-fluorimétrie

Fluoranthene		GC/FID

Pyrene		GC/FID

Benzo[a]anthracene		HPLC-fluorimétrie

Chrysene		GC/FID

Benzo[b]fluoranthene		GC/FID

Benzo[k]fluoranthene		GC/FID

Benzo[a]pyrene		HPLC-fluorimétrie

Indeno[1,2,3-cd]pyrene		GC/FID

Dibenzo[a,h]anthracene	HPLC-fluorimétrie

Benzo[ghi]perylene		GC/FID

Biphenyl		GC/FID

Perylene		GC/FID

Benzo[e]pyrene		HPLC-fluorimétrie

Uncertainty: 15 %

Chlorinated compounds:

GC/ECD on HP-5 and CP-SIL 19CB.

Uncertainty: 15 %

Impurity overcoat -chlordane on HP-5 and on CP-SIL 19CB -chlordane is overcoated by p,p’-DDE.

Lab 110

PAHs:

The samples were evaporated to eliminate cyclohexane and the residue was dissolved in acetonitrile. The extracts obtained were analysed by HPLC with fluorescence detector.

Chlorinated compounds:

GC/MS

Lab 113

Brief description of the method used for PAH:

1) Cleanup

None.

3 different dilutions in acetonitrile/water were prepared: 1:50, 1:100, 1:500.

2) Calibration

Seven point calibration with external standards in the range 0,2- 10 pg/µl using a linear calibration function.

3) Analysis

High Pressure Liquid Chromatography

HPLC System Hewlett Packard HP1100 with:

- Binary Gradient Pump G1312A;

- Autosampler G1313A;

- Vacuum-Degasser G1322A,

- Column Thermostat G1316A;

- Fluoreszensdetector G1321A;

- Chromatography Software (HP ChemStation for LC 3D Rev. A.08.03 [847]);

- HPLC column:	SEPSERV 250 mm * 2 mm Ultrasep ES PAK 6 µm;

- Precolumn: 		10 mm * 2 mm Ultrasep ES PAK 6 µm.

Injection volume:	15 µl

Gradient:

Time in min	% Acetonitril	% Water

0 -5	55	45

5- 30	55 to 100	45 auf 0

30 - 44	100	0

44,01	100 - 55	0 - 45

44,01 - 60	55	45

Wavelength programme:

Time in min 	Exitation	Emission	PMT

0,00	275	335	16

13,30	260	369	16

17,40	235	460	16

19,55	250	390	16

22,40	260	400	16

29,00	260	450	15

31,30	290	420	16

34,50	290	420	16

37,80	250	500	16

40,00	270	335	16

4) Limit of detection:

The limits of detection are roughly estimated for pure standard solutions from the calibration functions.

The estimate for Indeno(1,2,3-cd)pyrene is about 0,3 pg/µl and for all other PAH about 0,1 pg/µl.

5) Uncertainty

From our routinely performed analysis of PAH in rainwater we estimate the uncertainty at the limit of determination to about 40% and in the middle of the calibration range to about 20%.

Brief description of the method used for chlorinated POPs

1) Cleanup

A spatula of sodium sulphate is added on top of a commercially available silica gel column (Chromabond/Mackerey-Nagel) containing 500 mg of silica gel. The column is conditioned by elution with 3 ml acetone (3x), vacuum drying for 5 minutes, elution with 3 ml solvent mixture (70% n-hexane/30% toluene v/v).

500 µl of the extract (sample) and the internal standards -HCH and PCB 209 (10pg/µl) are passed through the conditioned silica gel column. Followed by elution with 7 X 2 ml solvent mixture.

The eluat is collected in a 15 ml centrifuge tube.

The eluat is concentrated at 35° C in a Turbo-Vap Evaporator to about 150 µl.

2) Calibration

Six point calibration in the range 0,5-10 pg/µl with two internal standards using a second order calibration function.

 -HCH (IS II) used as internal standard for HCB, -HCH, -HCH, p,p‘-DDE and PCB 101.

PCB 209 (IS I) used for p,p’-DDT, PCB 118, PCB 138, PCB 153 and PCB 180.

3) Analysis

Gas chromatography

Instrument:	HP 6890 Series	Agilent Technologies

Column:	RTX-CLP 	Restek

30 m x 0,25 mm,  0,25 µm

Carrier gas: 	Helium

Pressure:	1,33 bar

Flow rate:	1,7 ml/min (constant)

Oven programme:

Start temperature:	115 °C

Time (isotherm):	2 min

Heating rate 1:	4 °C/min	Ramp temperature 1: 260 °C

Time(isotherm):	1 min

Heating rate 2:	30 °C/min	Ramp temperature 2: 320 °C

Time (isotherm):	6 min

Mass spectrometry:

Instrument:	MSD 5973	Agilent Technologies

Ionisation technique:	NCI

Reactant gas: 	Methane

Source pressure: 	2,4 x 10 –4 torr

Ion source temperature: 	120°C

Electron energy: 	179.4 eV

Emission current: 	49,4 µA

Multiplier voltage:  	1895 V

Acquisition mode:	Selected ion monitoring (SIM)

SIM-program:

| Function | Time window (min) | Mass (amu) |  | Compound |
| --- | --- | --- | --- | --- |
|  |  | Quantification | Qualifier |  |
| 1 | 10,00-15,30 | 71 71 283,9 | 255 255 281,9 | -HCH -HCH HCB |
| 2 | 15,30-19,00 | 71 221 292 | 255 257 257 | -HCH PCB 28 PCB 52 |
| 3 | 19,00-24,75 | 325,9 318 325,9 | 327,9 316 327,9 | PCB 101 p, p‘-DDE PCB 118 |
| 4 | 24,75-28,00 | 359,8 359,8 71 | 361,8 361,8 248 | PCB 153 PCB 138 p, p‘-DDT |
| 5 | 28,00-36,00 | 393,8 499,7 | 395,8 497,7 | PCB 180 PCB 209 |

4) Limits of determination and detection

The limit of determination and the limit of detection are estimated for pure standard solutions (using 500 µl) from the calibration functions.

| Compound | Limit of determination in pg/µl | Limit of detection in pg/µl |
| --- | --- | --- |
| P, p' DDT | 0,5 | 0,17 |
| P, p' DDE | 0,5 | 0,17 |
| -HCH | 0,1 | 0,03 |
| HCB | 0,03 | 0,01 |
| -HCH | 0,1 | 0,03 |
| PCB 28 | 5 | 1,7 |
| PCB 52 | 5 | 1,7 |
| PCB 101 | 0,1 | 0,3 |
| PCB 118 | 0,05 | 0,02 |
| PCB 138 | 0,05 | 0,02 |
| PCB 153 | 0,03 | 0,01 |
| PCB 180 | 0,05 | 0,02 |

5) Uncertainty

From our routinely performed analysis of chlorinated compounds in human blood we estimate the uncertainty at the limit of determination to about 40-50% and in the middle of the calibration range to about 25%.

Lab 115

Chlorinated compounds:

The following procedure has been applied to the samples.

- an internal standard of 2.5 ng PCB 6, PCB 65, PCB 207 and e-HCH was added

- -two subsamples of 25 perc. were analyzed

- a HPLC fractionation had been performed with the subsamples which led to two fractions :

1. fraction: PCBs, HCB, pp-DDE

2. fraction: Chlordanes, HCHs, pp-DDT

- the samples were measured using GC-MS in the MSMS mode. ( CP-SIL 8 column, varian saturn 2000 GC-MS)

Detection limits are between 500 fg-1 pg for the target compounds. The analytical precision is in the order of 15 perc. standard deviation. pp-DDT was below the detection limit.

Lab 116

PAHs

Method 1:	HPLC in adaption to ISO/CD 16362

No clean-up

Calibration one point 50 pg/µl

Limits of determination:	ng/m3

Naphthalene, Acenaphthene	0,04

Acenaphthylene	0,08

all others each	0,02

Uncertainty: +/- 25%

Method 2:	GC-MS in adaption to DIN ISO 12884, SIM-Mode

No clean-up

Calibration 3-point 100 pg/µl to 10 ng/µl

Limits of determination:	ng/m3

all components 	0,04

Uncertainty: +/- 25%

Method 1 + 2 comparatively measured against NIST SRM 1647d

Chlorinated compounds:

Method: GC-ECD two-column-technique

in adaption of DIN 38409-F2

no clean-up

Calibration 6-point 10 - 100 pg/µl

Limits of determination: 	pg/m3

all components 	10

Uncertainty: +/- 25%

Lab 118

Chlorinated compounds:

In round 1 the solution was diluted with internal standard (TCN) and analysed by GC-ECD using DB5 and DB1701 (JW-Scientific), both 60m, 0.25 mm i.d., 0.25 µm film.

The estimated uncertainty is close to 10%.

In round 2 the samples were cleaned up with conc. sulphuric acid, evaporated about 50% and redissolved in isooctane containing TCN (ISTD).

Lab 120

Chlorinated compounds:

GC/MS analysis (round 1 only).

Using working standards at 0.06/0.1 ug/ml and 0.03/0.05 ug/ml for the chlorinated pesticides and CB congeners respectively, with SIR detection.

The results are described as very crude, because of instrument difficulties.

Lab 121

Chlorinated compounds:

Method:	GC

Detector:	ECD

Column:	PE 608, length - 30 m, ID - 0,53 mm.

Lab 122

PAHs:

Method of analysis round 1

0.2 mL of the "unknown" mixture was fractionated on a thin layer of Al2O3 with solvent mixtures of petroleum ether: diethyl ether in the ratio 40:1. PAHS were divided into 4 fractions.  Each fraction was eluted by 50 mL of n-hexane. Chromatographic eluates were evaporated to 10 mL. Different PAHs were determined in separate fractions with different Rf. Identification and quantity determination of PAHs was performed by the spectrofluorescent method with luminescent spectrometer СДЛ-2 "ЛОМО" using the method of standard addition.

Quantitative analysis of BP was performed by the spectrofluorescent method in liquid nitrogen at 77K. Excitation of BP was performed at 298 nm and fluorescence was recorded at 403 nm. The fluorescent analysis was performed by a diffraction spectrometer (DФС-12 “ЛOMO”). The sensitivity of the method: 0.1 ng mL.

The characteristics of analytical method are presented in Table A1.1.

Table A1.1:	Characteristics of analytical method.

| PAHs Compound | exc., nm | fl., nm | Detection limit, g.mL |
| --- | --- | --- | --- |
| 1. Phenanthrene | 255 | 347 | 0.1 |
| 2. Pyrene | 330 | 372 | 0.01 |
| 3. Chrysene | 300 | 362 | 0.01 |
| 4. Benz(a)Anthracene | 280 | 385 | 0.1 |
| 5. Benzo(a)pyrene | 298 | 403 | 0.0001 |
| 6. Benz(g,h,i)Perylene | 300 | 420 | 0.1 |
| 7. Dibenz(a,h)Anthracene | 300 | 394 | 0.01 |

Method of analysis round 2, B(a)P only

0.2 mL of each extract was fractionated on a thin layer of Al2O3 with solvent mixtures of hexene: benzene 6:1. All compounds were divided into 3 fractions. Fraction of BaP was eluted by 50 mL of n-hexane. Chromatographic eluates of BaP fraction were evaporated to 5 mL. Quantitative analysis of BaP was performed by the spectrofluorescent method in liquid nitrogen at 77K using the method of standard addition. Excitation of BP was performed at 360 nm and fluorescence was recorded at 403 nm. The fluorescent analysis was performed by a diffraction spectrometer (DФС-12 “ЛOMO”). The sensitivity of the method: 0.1 ng mL.

Lab 123

PAHs

Instrument:	Gas Chromatograph HP 5890.

Column:	Supelco PTE-5 30 m x 0.25 mm, 0.25 m

Carrier:	Hydrogen 1ml/min (constant flow)

Make up gas: 	Nitrogen 35 ml/min

Hydrogen: 	30 ml/min

Air: 	360 ml/min

Injector: 	Splitless at 280C

Detector: 	FID at 325ºC

Temperature program: 	70ºC, 1 min

8º/min

310ºC, 3 min

Calibration standard: 	NILU solution

Chlorinated compounds:

Instrument: 	Gas Chromatograph HP 5890.

Column: 	Supelco PTE-5 30 m x 0.25 mm, 0.25 m

Carrier: 	Hydrogen 1 ml/min (constant flow)

Make up gas: 	Nitrogen 55 ml/min

Injector: 	Splitless at 280C

Detector: 	ECD at 300ºC

Temperature program: 	120ºC, 1 min

30º/min to 180ºC

2º/min to 220ºC

25º/min to 300ºC

300ºC, 8 min

Calibration standard: 	NILU solution

Lab 129A

PAHs: Round 1

The vials were opened and the content was transferred to screw capped vials and weighed. Sub-samples of each solution were spiked with QS internal standards (1000 ng each of D8-naphtalene, D8-acenaphtylene, D10-acenaphtene, D10-anthracene, D10-pyrene, D12-benzo(b)fluoranthene, D14-dibenzo(ah)anthracene), after the addition of RS standard (1000 ng of each of D10-fluorene, D12-chrysene and D12-benzo(ghi)perylene) the 1 ml sample (cleaned up by column chromatography if necessary) was measured by GC/MSD for PAH after the addition of a syringe standard (1000 ng of each of D10-fluoranthene and D12-benzo(a)pyrene)

Measurements were achieved using a HP 5970 GC/MS equipped with a 60m ZB5 0.2 m film capillary column. The MS was operated in SIM mode.

GC conditions:

Injector temp	280C

Initial temp	60C for 0.5 min

First ramp	20C/min to 240C

Second ramp	3.5C/min to 280C

Third ramp	5C/min to 300C

Final temp	300C for 24 min

Interface temp	275C

PAHs round 2:

As described above, but the whole extract and clean-up were used:

Each extract was subjected to a normal phase column chromatographic cleanup using alumina and silica. The collected PAH fraction was then concentrated and spiked with a further mixture containing deuterated benzo(a)pyrene and fluoranthene to facilitate measurement of recoveries of the internal standards used for quantitation.

The total amount of each PAH was measured in the clean samples using isotope dilution GC/MS.

Chlorinated compounds:

Round 1

The vials were opened and the content was transferred to screw capped vials and weighed. Sub-samples of each solution were spiked with QS internal standards (approximately 100 ng each of 13C-labelled hexachlorbenzene, PCB-28, p,p’-DDE, DDT, PCB-180)

Measurements were achieved using a HP 5970 GC/MS equipped with a 60m ZB5 0.2m film capillary column. The MS was operated in SIM mode.

GC conditions:

Injector temp	300C

Initial temp	60C for 10 min

First ramp	10C/min to 180C

Second ramp	5C/min to 240C

Third ramp	8C/min to 280C

Final temp	280C for 12 min

Interface temp	280C

Round 2

As above, but the whole extract was used.

Lab 129L

PAHs: Round 1 only

The unknown EMEP PAH solution was analysed on HPLC and GC-MS, and quantified using a set of calibration standards made up from the known EMEP solution and using Lancasters calibration standards. The standards used by Lancaster for HPLC quantification and GC-MS quantification were bought at different times, and prepared by different people. For the HPLC analysis two dilutions of the unknown solution were prepared, dilution A was run in duplicate and dilution B in triplicate. A single dilution was run for GC-MS in triplicate.

Chlorinated compounds: Round 1 only

Two separate dilutions (A and B) of the EMEP unknown solution were analysed in duplicate using GC-MS, and quantified using the known EMEP solution and Lancasters calibration standards.

Lab 130

PAHs

Varian 3400 GC equipped with Split/Splitless injector operated at 2660. Column Supelco PTE-5 30m, 0.32mm id, 0.25m film. Carrier gas hydrogen, 1ml/min measured at 2100. Column temperature was linearly programmed from 1110 to 2900 at 4.30/min. Detector FID 3100.

Confirmation on:

Ion Trap Detector ITD-705 Finnigan was used with Varian 3400 GC equipped with Split/Splitless injector operated at 2660. Column Supelco PTE-5 30m, 0.25mm id, 0.25m film, inserted directly in Ion Trap via transfer line at 2400. Carrier gas hydrogen, 1ml/min measured at 2100. Ion manifold and exit nozzle temperatures of 2400 were used. Column temperature was linearly programmed from 600 to 2850 at 4.30/min. Scan range was 39-333 Daltons, 1 scan/second (5 micro scans were averaged). Tuning parameters were: 0,50,50,100; AGC on, Background mass 33. Version 3 of ITDS software was used.

Lab 131

PAHs: Round 2

To each extract was added an internal standard containing D10-2-methylnaphtalene, D10-acenaphtene, D10-anthracene, D10-pyrene, D12-benz(a)anthracene, D12-benzo(e)pyrene and D12-benzo(ghi)perylene.

The extract was cleaned up, first using liquid/liquid partitioning between cyclohexane and dimethyl formamide-water.

Secondly the sample was subjected to a normal phase column chromatographic clean-up using silica.

After pre-concentration to about 100l a recovery standard containing D8-biphenyl, D10-fluoranthene and D12-perylene was added before quantification using HP 5973 GC/MS.

Gas chromatography

Instrument: 	HP 6890 Series	Agilent

Column:	CP-SIL 8CB 	Varian

25 m x 0.25 mm, 0.12 µm

Carrier gas: 	Helium

Pressure:	12.6 psi

1l sample was injected splitless.

Temperature programme:

Start temperature:	50ºC

Time (isotherm):	1 min

Heating rate 1:	20°C/min	Ramp temperature 1: 100°C

Heating rate 2:	  5°C/min	Ramp temperature 2: 300°C

Time (isotherm):	3.5 min

Injector temperature:	300°C

Interface temperature:	310°C

The MS was operated in SIM mode recording the molecular mass of the PAHs measured.

Uncertainty: 	+/- 25%

Limits of detection (S/N 3:1): 	2-10 pg/m3

Chlorinated compounds: Round 2

To each extract was added an internal standard containing 13C labeled -HCH, HCB, PCB-28, PCB-52, PCB-101, PCB-118, PCB-153, PCB-180, p,p’-DDE, p,p’-DDT and 13C-6D--HCH.

The extracts were treated with concentrated sulphuric acid to oxidize non-persistent compounds before normal phase column chromatographic clean-up using silica.

After pre-concentration to about 100l a recovery standard containing tetrachloronaphtalene.

The HCHs and the chlordanes were analysed using a HP 5989 “ENGINE” GC/MS-system and negative ion chemical ionization.

The column used was a 25m x 0.2mm 0.11m Ultra-2 from J&W.

1l sample was injected splitless.

Uncertainty: 	+/- 25%

Limits of detection (S/N 3:1): 	0.05 pg/m3

DDE and DDT were analysed using a VG AutoSpec with electron impact ionization and mass resolution 10.000.

The column used was a 25m x 0.2mm 0.33m HP-1 from J&W.

1l sample was injected splitless.

Uncertainty: 	+/- 25%

Limits of detection (S/N 3:1): 	0.05 pg/m3

The PCBs were analysed using a VG AutoSpec with electron impact ionization and mass resolution 10.000.

The column used was a 50m x 0.22mm 0.11m HT-8 from SGE.

1l sample was injected splitless.

Uncertainty: 	+/- 25%

Limits of detection (S/N 3:1): 	0.05 pg/m3

- Organochlorine compounds, round 1

On the following pages the results from round 1 are shown graphically. The data from each compound is shown as deviation from expected value (upper graph) and as concentration found in the sample (lower graph).

	GC/ECD

	GC/MS

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | pp'-DDT |  |  |
| --- | --- | --- | --- | --- |
| Theoretical value: |  | 20,0 |  |  |
| Unit: |  | pg/ul |  |  |
| Number of laboratories: |  | 14 |  |  |
| Aritm. Mean |  | 21,7 |  |  |
| Median |  | 21,9 |  |  |
| Lab no. | pg/ul |  | Lab no. | pg/ul |
| 105 | 34,4 |  | 102 | 21,8 |
| 121 | 27,0 |  | 103 | 20,2 |
| 104 | 26,7 |  | 113 | 19,7 |
| 123 | 24,7 |  | 101 | 19,6 |
| 110 | 24,0 |  | 118 | 18,3 |
| 108 | 22,3 |  | 129L | 15,9 |
| 129A | 22,0 |  | 109 | 6,8* |
| *Not used for determination of mean |  |  |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | pp'-DDE |
| --- | --- | --- |
| Theoretical value: |  | 40,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 40,9 |
| Median |  | 40,1 |
| Lab no. | pg/ul |  |
| 105 | 66,4* |  |
| 123 | 48,4 |  |
| 104 | 47,9 |  |
| 102 | 47,3 |  |
| 129L | 44,4 |  |
| 113 | 40,5 |  |
| 103 | 40,1 |  |
| 101 | 39,0 |  |
| 118 | 38,7 |  |
| 129A | 38,0 |  |
| 108 | 34,7 |  |
| 121 | 28,0 |  |
| 109 | 18,7* |  |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | g-Chlordane |
| --- | --- | --- |
| Theoretical value: |  | 7,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 12 |
| Aritm. Mean |  | 12,2 |
| Median |  | 10,7 |
| Lab. No. | pg/ul |  |
| 123 | 220,0 | * |
| 109 | 34,7 |  |
| 129A | 13,0 |  |
| 104 | 11,7 |  |
| 102 | 11,0 |  |
| 103 | 10,7 |  |
| 105 | 10,6 |  |
| 101 | 9,8 |  |
| 129L | 8,7 |  |
| 121 | 8,5 |  |
| 118 | 8,1 |  |
| 108 | 7,3 |  |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | a-Chlordane |
| --- | --- | --- |
| Theoretical value: |  | 15,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 12 |
| Aritm. Mean |  | 15,9 |
| Median |  | 14,9 |
| Lab. No. | pg/ul |  |
| 123 | 21,4 |  |
| 102 | 19,7 |  |
| 104 | 17,9 |  |
| 103 | 17,3 |  |
| 129A | 16,0 |  |
| 109 | 15,0 |  |
| 129L | 14,8 |  |
| 101 | 14,6 |  |
| 118 | 13,9 |  |
| 105 | 13,3 |  |
| 108 | 11,0 |  |
| 121 | 6,0 | * |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | g-HCH |
| --- | --- | --- |
| Theoretical value: |  | 300 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 275 |
| Median |  | 284 |
| Lab. No. | pg/ul |  |
| 120 | 2300 | * |
| 101 | 328 |  |
| 102 | 325 |  |
| 129L | 323 |  |
| 103 | 295 |  |
| 110 | 292 |  |
| 118 | 291 |  |
| 123 | 284 |  |
| 113 | 284 |  |
| 108 | 279 |  |
| 121 | 275 |  |
| 109 | 274 |  |
| 104 | 272 |  |
| 129A | 230 |  |
| 105 | 104* |  |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | a-HCH |
| --- | --- | --- |
| Theoretical value: |  | 120 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 106 |
| Median |  | 115 |
| Lab. No. | pg/ul |  |
| 129A | 800* |  |
| 129L | 138 |  |
| 102 | 130 |  |
| 123 | 128 |  |
| 110 | 119 |  |
| 101 | 118 |  |
| 113 | 117 |  |
| 104 | 112 |  |
| 118 | 111 |  |
| 108 | 111 |  |
| 109 | 104 |  |
| 103 | 104 |  |
| 121 | 58 |  |
| 105 | 34 |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | HCB |
| --- | --- | --- |
| Theoretical value: |  | 150 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 143 |
| Median |  | 142 |
| Lab. No. | pg/ul |  |
| 129L | 175 |  |
| 102 | 156 |  |
| 108 | 155 |  |
| 118 | 155 |  |
| 103 | 151 |  |
| 109 | 144 |  |
| 121 | 142 |  |
| 110 | 141 |  |
| 101 | 137 |  |
| 104 | 130 |  |
| 113 | 128 |  |
| 123 | 124 |  |
| 129A | 120 |  |
| 105 | 91 | * |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | PCB 28 |
| --- | --- | --- |
| Theoretical value: |  | 36,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 36,5 |
| Median |  | 37,9 |
| Lab. No. | pg/ul |  |
| 113 | 55,5 | * |
| 129L | 46,2 |  |
| 105 | 45,3 |  |
| 102 | 43,1 |  |
| 104 | 39,6 |  |
| 108 | 38,4 |  |
| 123 | 38,2 |  |
| 121 | 37,5 |  |
| 101 | 36,3 |  |
| 118 | 34,0 |  |
| 110 | 33,7 |  |
| 109 | 31,0 |  |
| 129A | 29,0 |  |
| 103 | 22,6 |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | PCB 52 |  |  |
| --- | --- | --- | --- | --- |
| Theoretical value: |  | 76,0 |  |  |
| Unit: |  | pg/ul |  |  |
| Number of laboratories: |  | 15 |  |  |
| Aritm. Mean |  | 70,0 |  |  |
| Median |  | 68,5 |  |  |
| Lab. No. | pg/ul |  | Lab. No. | pg/ul |
| 129L | 78,9 |  | 108 | 68,3 |
| 118 | 78,9 |  | 121 | 68,0 |
| 113 | 77,4 |  | 110 | 66,0 |
| 102 | 77,3 |  | 103 | 65,7 |
| 101 | 71,8 |  | 109 | 64,0 |
| 123 | 71,3 |  | 129A | 63,0 |
| 105 | 68,6 |  | 104 | 60,7 |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | PCB 101 |
| --- | --- | --- |
| Theoretical value: |  | 42,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 40,3 |
| Median |  | 38,2 |
| Lab. no. | pg/ul |  |
| 105 | 56,6 |  |
| 109 | 45,7 |  |
| 129L | 45,0 |  |
| 108 | 43,6 |  |
| 101 | 39,6 |  |
| 113 | 39,3 |  |
| 102 | 38,6 |  |
| 104 | 37,8 |  |
| 118 | 37,8 |  |
| 110 | 37,3 |  |
| 129A | 33,0 |  |
| 103 | 29,0 |  |
| 121 | 7,6* |  |
| 123 | 2,8* |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | PCB 118 |
| --- | --- | --- |
| Theoretical value: |  | 53,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 52,0 |
| Median |  | 53,0 |
| Lab. no. | pg/ul |  |
| 105 | 108,5* |  |
| 109 | 65,7 |  |
| 102 | 64,8 |  |
| 129L | 62,9 |  |
| 108 | 57,1 |  |
| 118 | 53,9 |  |
| 104 | 53,5 |  |
| 123 | 52,4 |  |
| 101 | 51,3 |  |
| 110 | 50,7 |  |
| 113 | 49,4 |  |
| 129A | 44,0 |  |
| 121 | 36,7 |  |
| 103 | 34,1 |  |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | PCB 138 |
| --- | --- | --- |
| Theoretical value: |  | 24,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 24,2 |
| Median |  | 24,6 |
| Lab. no. | pg/ul |  |
| 105 | 36,8 |  |
| 102 | 28,4 |  |
| 129L | 27,2 |  |
| 109 | 27,0 |  |
| 123 | 26,5 |  |
| 108 | 25,2 |  |
| 104 | 24,8 |  |
| 110 | 24,3 |  |
| 113 | 22,9 |  |
| 118 | 22,3 |  |
| 101 | 21,9 |  |
| 129A | 16,0 |  |
| 103 | 11,4 |  |
| 121 | 6,0* |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | PCB 153 |
| --- | --- | --- |
| Theoretical value: |  | 22,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 20,9 |
| Median |  | 21,2 |
| Lab. no. | pg/ul |  |
| 105 | 32,1* |  |
| 102 | 27,5 |  |
| 123 | 25,7 |  |
| 108 | 23,2 |  |
| 129L | 23,1 |  |
| 113 | 21,8 |  |
| 104 | 21,4 |  |
| 110 | 21,0 |  |
| 118 | 20,5 |  |
| 101 | 20,4 |  |
| 129A | 20,0 |  |
| 109 | 17,7 |  |
| 103 | 15,8 |  |
| 121 | 13,9 |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/l.

| Compound: |  | PCB 180 |
| --- | --- | --- |
| Theoretical value: |  | 30,0 |
| Unit: |  | pg/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 29,3 |
| Median |  | 29,5 |
| Lab. no. | pg/ul |  |
| 105 | 35,9 |  |
| 123 | 33,4 |  |
| 102 | 33,1 |  |
| 118 | 30,6 |  |
| 113 | 30,3 |  |
| 129L | 29,6 |  |
| 101 | 29,6 |  |
| 108 | 29,4 |  |
| 109 | 28,7 |  |
| 104 | 28,6 |  |
| 110 | 27,7 |  |
| 129A | 25,0 |  |
| 121 | 19,7* |  |
| 103 | 17,9* |  |
| *Not used for determination of mean |  |  |

- PAH compounds, round 1

On the following pages the results from round 1 are shown graphically. The data from each compound is shown as deviation from expected value (upper graph) and as concentration found in the sample (lower graph).

	HPLC

	GC/MS

GC/FID

	TLC

	HPLC and GC/MS

	GC/FID and GC/MS

Results obtained by each laboratory.

Upper figure: Deviation from median in %.. Lower figure: Concentrations in ng/l.

| Compound: |  | Naphtalene |
| --- | --- | --- |
| Theoretical value: |  | 8,33 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 12 |
| Aritm. Mean |  | 8,119 |
| Median |  | 8,174 |
| Lab. no. | ng/ul |  |
| 105 GC/MS | 10,610 |  |
| 101 | 10,330 |  |
| 105 HPLC | 8,550 |  |
| 109 | 8,533 |  |
| 129A | 8,253 |  |
| 107 | 8,190 |  |
| 123 | 8,158 |  |
| 108 | 7,988 |  |
| 129L GC/MS | 7,520 |  |
| 129L HPLC | 5,820 |  |
| 104 | 5,353 |  |
| 102 | 1,580 | * |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Acenaphtylene |
| --- | --- | --- |
| Theoretical value: |  | 2,00 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 12 |
| Aritm. Mean |  | 2,735 |
| Median |  | 2,680 |
| Lab. no. | ng/ul |  |
| 105GC/MS | 3,570 |  |
| 101 | 3,190 |  |
| 105 HPLC | 2,990 |  |
| 109 | 2,967 |  |
| 129A | 2,863 |  |
| 107 | 2,680 |  |
| 108 | 2,600 |  |
| 123 | 2,577 |  |
| 129L GC/MS | 2,493 |  |
| 103 | 2,477 |  |
| 104 | 1,680 |  |
| 102 | 0,222 | * |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Acenaphthene |
| --- | --- | --- |
| Theoretical value: |  | 4,17 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 12 |
| Aritm. Mean |  | 4,04 |
| Median |  | 3,99 |
| Lab. no. | ng/ul |  |
| 123 | 5,32 |  |
| 105 GC/MS | 5,15 |  |
| 101 | 4,58 |  |
| 105 HPLC | 4,10 |  |
| 109 | 4,07 |  |
| 107 | 4,05 |  |
| 129A | 3,93 |  |
| 129L GC/MS | 3,57 |  |
| 129L HPLC | 3,57 |  |
| 108 | 3,55 |  |
| 104 | 2,53 |  |
| 102 | 0,24 | * |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Fluorene |
| --- | --- | --- |
| Theoretical value: |  | 4,00 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 3,75 |
| Median |  | 3,68 |
| Lab. no. | ng/ul |  |
| 123 | 4,86 |  |
| 105 GC/MS | 4,77 |  |
| 101 | 4,17 |  |
| 105 HPLC | 4,13 |  |
| 130 | 4,10 |  |
| 109 | 3,80 |  |
| 129A | 3,69 |  |
| 107 | 3,67 |  |
| 129L HPLC | 3,59 |  |
| 108 | 3,58 |  |
| 129L GC/MS | 3,23 |  |
| 103 | 3,02 |  |
| 104 | 2,16 |  |
| 102 | 1,30 | * |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Phenanthrene |
| --- | --- | --- |
| Theoretical value: |  | 8,33 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 15 |
| Aritm. Mean |  | 7,601 |
| Median |  | 7,890 |
| Lab. no. | ng/ul |  |
| 105 GC/MS | 10,600 |  |
| 123 | 9,455 |  |
| 101 | 9,353 |  |
| 105 HPLC | 8,680 |  |
| 129A | 8,413 |  |
| 109 | 8,400 |  |
| 108 | 8,011 |  |
| 107 | 7,890 |  |
| 129L HPLC | 7,846 |  |
| 103 | 7,513 |  |
| 129L GC/MS | 7,147 |  |
| 122 | 5,700 |  |
| 102 | 5,280 |  |
| 104 | 4,983 |  |
| 130 | 4,750 |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Anthracene |
| --- | --- | --- |
| Theoretical value: |  | 0,42 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 0,392 |
| Median |  | 0,354 |
| Lab. no. | ng/ul |  |
| 101 | 0,520 |  |
| 129A | 0,520 |  |
| 105 GC/MS | 0,510 |  |
| 123 | 0,464 |  |
| 129L GC/MS | 0,403 |  |
| 105 HPLC | 0,380 |  |
| 108 | 0,354 |  |
| 109 | 0,353 |  |
| 129L HPLC | 0,348 |  |
| 107 | 0,336 |  |
| 110 | 0,320 |  |
| 103 | 0,314 |  |
| 102 | 0,274 |  |
| 104 | 0,190 | * |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Fluoranthene |
| --- | --- | --- |
| Theoretical value: |  | 4,17 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 15 |
| Aritm. Mean |  | 3,735 |
| Median |  | 3,880 |
| Lab. no. | ng/ul |  |
| 105 GC/MS | 5,020 |  |
| 123 | 4,328 |  |
| 101 | 4,237 |  |
| 129L HPLC | 4,234 |  |
| 105 HPLC | 4,160 |  |
| 109 | 4,067 |  |
| 108 | 3,895 |  |
| 129A | 3,880 |  |
| 107 | 3,855 |  |
| 110 | 3,703 |  |
| 103 | 3,516 |  |
| 129L GC/MS | 3,500 |  |
| 102 | 2,700 |  |
| 130 | 2,550 |  |
| 104 | 2,373 |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Pyrene |
| --- | --- | --- |
| Theoretical value: |  | 3,33 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 16 |
| Aritm. Mean |  | 3,10 |
| Median |  | 3,30 |
| Lab. no. | ng/ul |  |
| 105 GC/MS | 4,13 |  |
| 101 | 3,59 |  |
| 129L HPLC | 3,49 |  |
| 105 HPLC | 3,45 |  |
| 109 | 3,43 |  |
| 110 | 3,41 |  |
| 123 | 3,40 |  |
| 108 | 3,35 |  |
| 129A | 3,24 |  |
| 103 | 3,21 |  |
| 107 | 3,07 |  |
| 129L GC/MS | 2,72 |  |
| 102 | 2,32 |  |
| 104 | 1,88 |  |
| 130 | 1,80 |  |
| 122 | 1,20 | * |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Benzo[a]anthracene |
| --- | --- | --- |
| Theoretical value: |  | 0,58 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 16 |
| Aritm. Mean |  | 0,558 |
| Median |  | 0,533 |
| Lab. no. | ng/ul |  |
| 122 | 2,000 | * |
| 123 | 1,377 |  |
| 105 GC/MS | 0,650 |  |
| 105 HPLC | 0,630 |  |
| 129L HPLC | 0,576 |  |
| 110 | 0,557 |  |
| 109 | 0,547 |  |
| 103 | 0,535 |  |
| 107 | 0,531 |  |
| 129L GC/MS | 0,517 |  |
| 102 | 0,516 |  |
| 129A | 0,503 |  |
| 101 | 0,423 |  |
| 108 | 0,396 |  |
| 130 | 0,350 |  |
| 104 | 0,267 |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Chrysene |  |  |
| --- | --- | --- | --- | --- |
| Theoretical value: |  | 1,67 |  |  |
| Unit: |  | ng/ul |  |  |
| Number of laboratories: |  | 16 |  |  |
| Aritm. Mean |  | 1,28 |  |  |
| Median |  | 1,35 |  |  |
| Lab. no. | ng/ul |  | Lab. no. | ng/ul |
| 122 | 13,50* |  | 129A | 1,35 |
| 102 | 1,70 |  | 103 | 1,28 |
| 105 HPLC | 1,66 |  | 129L GC/MS | 1,20 |
| 107 | 1,58 |  | 108 | 1,17 |
| 109 | 1,53 |  | 101 | 1,13 |
| 105 GC/MS | 1,47 |  | 130 | 0,95 |
| 129L HPLC | 1,43 |  | 123 | 0,71 |
| 110 | 1,43 |  | 104 | 0,55 |
| *Not used for determination of mean |  |  |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Benzo[b]fluoranthene |
| --- | --- | --- |
| Theoretical value: |  | 2,18 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 13 |
| Aritm. Mean |  | 2,028 |
| Median |  | 2,080 |
| Lab. no. | ng/ul |  |
| 103 | 2,791 |  |
| 109 | 2,300 |  |
| 102 | 2,220 |  |
| 129L HPLC | 2,172 |  |
| 105 GC/MS | 2,170 |  |
| 110 | 2,093 |  |
| 105 HPLC | 2,080 |  |
| 129A | 2,023 |  |
| 107 | 1,850 |  |
| 129L GC/MS | 1,787 |  |
| 101 | 1,747 |  |
| 104 | 1,103 |  |
| 123 | 1,004 | * |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Benzo[k]fluoranthene |
| --- | --- | --- |
| Theoretical value: |  | 0,50 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 14 |
| Aritm. Mean |  | 0,487 |
| Median |  | 0,484 |
| Lab. no. | ng/ul |  |
| 123 | 1,377 | * |
| 105 GC/MS | 0,620 |  |
| 109 | 0,530 |  |
| 129L GC/MS | 0,530 |  |
| 105 HPLC | 0,520 |  |
| 101 | 0,513 |  |
| 129L HPLC | 0,492 |  |
| 103 | 0,476 |  |
| 110 | 0,473 |  |
| 102 | 0,462 |  |
| 129A | 0,453 |  |
| 108 | 0,442 |  |
| 104 | 0,417 |  |
| 107 | 0,400 |  |
| *Not used for |  |  |
| determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Benzo[a]pyrene |
| --- | --- | --- |
| Theoretical value: |  | 0,42 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 16 |
| Aritm. Mean |  | 0,355 |
| Median |  | 0,348 |
| Lab. no. | ng/ul |  |
| 123 | 1,079 | * |
| 122 | 0,990 | * |
| 105 GC/MS | 0,540 |  |
| 105 HPLC | 0,420 |  |
| 110 | 0,400 |  |
| 109 | 0,393 |  |
| 129L HPLC | 0,356 |  |
| 129L GC/MS | 0,353 |  |
| 129A | 0,350 |  |
| 103 | 0,347 |  |
| 108 | 0,338 |  |
| 101 | 0,330 |  |
| 102 | 0,314 |  |
| 107 | 0,305 |  |
| 104 | 0,267 |  |
| 130 | 0,250 |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Indeno[1,2,3-cd]pyrene |
| --- | --- | --- |
| Theoretical value: |  | 1,33 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 13 |
| Aritm. Mean |  | 1,159 |
| Median |  | 1,200 |
| Lab. no. | ng/ul |  |
| 123 | 1,351 |  |
| 109 | 1,333 |  |
| 103 | 1,324 |  |
| 102 | 1,240 |  |
| 110 | 1,227 |  |
| 107 | 1,210 |  |
| 105 HPLC | 1,200 |  |
| 108 | 1,152 |  |
| 105 GC/MS | 1,150 |  |
| 129A | 1,107 |  |
| 129L GC/MS | 1,053 |  |
| 101 | 0,887 |  |
| 104 | 0,837 |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Dibenzo[a,h]anthracene |
| --- | --- | --- |
| Theoretical value: |  | 0,17 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 15 |
| Aritm. Mean |  | 0,178 |
| Median |  | 0,160 |
| Lab. no. | ng/ul |  |
| 123 | 0,742 | * |
| 105 GC/MS | 0,460 |  |
| 129L GC/MS | 0,263 |  |
| 105 HPLC | 0,260 |  |
| 103 | 0,193 |  |
| 109 | 0,177 |  |
| 129A | 0,163 |  |
| 129L HPLC | 0,160 |  |
| 108 | 0,152 |  |
| 110 | 0,147 |  |
| 107 | 0,125 |  |
| 101 | 0,117 |  |
| 104 | 0,100 |  |
| 102 | 0,099 |  |
| 122 | 0,083 |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Benzo[ghi]perylene |
| --- | --- | --- |
| Theoretical value: |  | 1,00 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 15 |
| Aritm. Mean |  | 0,946 |
| Median |  | 0,937 |
| Lab. no. | ng/ul |  |
| 122 | 8,700 | * |
| 123 | 1,237 |  |
| 102 | 1,100 |  |
| 109 | 1,100 |  |
| 105 HPLC | 1,070 |  |
| 129L HPLC | 0,968 |  |
| 105 GC/MS | 0,950 |  |
| 110 | 0,937 |  |
| 107 | 0,935 |  |
| 108 | 0,919 |  |
| 101 | 0,887 |  |
| 129A | 0,857 |  |
| 103 | 0,812 |  |
| 129L GC/MS | 0,787 |  |
| 104 | 0,680 |  |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Biphenyl |
| --- | --- | --- |
| Theoretical value: |  | 7,98 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 8 |
| Aritm. Mean |  | 7,27 |
| Median |  | 6,88 |
| Lab. no. | ng/ul |  |
| 123 | 9,78 |  |
| 129A | 8,26 |  |
| 109 | 7,27 |  |
| 105 HPLC | 6,94 |  |
| 107 | 6,82 |  |
| 129L GC/MS | 6,58 |  |
| 104 | 5,21 |  |
| 105 GC/MS | 0,70 | * |
| *Not used for determination of mean |  |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Perylene |
| --- | --- | --- |
| Theoretical value: |  | 1,20 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 8 |
| Aritm. Mean |  | 1,12 |
| Median |  | 1,14 |
| Lab. no. | ng/ul |  |
| 109 | 1,43 |  |
| 107 | 1,31 |  |
| 129A | 1,19 |  |
| 105 HPLC | 1,17 |  |
| 130 | 1,10 |  |
| 129L GC/MS | 1,01 |  |
| 123 | 0,93 |  |
| 104 | 0,82 |  |

Results obtained by each laboratory.

Upper figure: Deviation from median in %. Lower figure: Concentrations in ng/l.

| Compound: |  | Benzo[e]pyrene |
| --- | --- | --- |
| Theoretical value: |  | 0,43 |
| Unit: |  | ng/ul |
| Number of laboratories: |  | 9 |
| Aritm. Mean |  | 0,440 |
| Median |  | 0,380 |
| Lab. no. | ng/ul |  |
| 123 | 1,156 | * |
| 103 | 0,806 |  |
| 105 HPLC | 0,490 |  |
| 105 GC/MS | 0,460 |  |
| 107 | 0,380 |  |
| 129A | 0,377 |  |
| 129L GC/MS | 0,360 |  |
| 104 | 0,333 |  |
| 109 | 0,317 |  |
| *Not used for determination of mean |  |  |

- Organochlorine compounds, round 2,
sample 1 and 2

The following pages present the results from round 2 graphically. The results for each compound are shown on two pages with the left-hand page giving the sample 1 results and the right-hand side the corresponding data from sample 2.

The results have been given as deviations from the median in per cent (upper graph) and as concentration based on a 500 m3 sample volume (lower graph).

	GC/ECD

	GC/MS

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories:14	Median 0,900

|  | Dev. f median |  |
| --- | --- | --- |
| Lab. no. | % | pg/m3 |
| 129A | 4989 | 45,8 |
| 121 | 311 | 3,7 |
| 105 | 67 | 1,5 |
| 113 | 44 | 1,292 |
| 131 | 0 | 0,9 |
| 118 | -2 | 0,88 |
| 103 | -3 | 0,869333 |
| 131 | -11 | 0,8 |
| 104 | -67 | 0,3 |
| 108 |  | <43,82 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | <5 |
| 109 |  | <2,9 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median	1,17

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 129A | 10075 | 119,3 |
| 121 | 241 | 4 |
| 105 | 62 | 1,9 |
| 118 | 31 | 1,54 |
| 113 | 11 | 1,296 |
| 103 | -11 | 1,0489333 |
| 131 | -32 | 0,8 |
| 131 | -32 | 0,8 |
| 104 | -83 | 0,2 |
| 123 | -83 | 0,2 |
| 108 |  | <44,442 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | < 5 |
| 109 |  | <2,9 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15 	Median	1,7007

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 113 | 297 | 6,749 |
| 123 | 59 | 2,7 |
| 131 | 18 | 2 |
| 131 | 12 | 1,9 |
| 103 | 0 | 1,701333333 |
| 115 | 0 | 1,7 |
| 118 | -5 | 1,62 |
| 104 | -9 | 1,545454545 |
| 105 | -12 | 1,5 |
| 109 | -18 | 1,4 |
| 108 |  | <175,28 |
| 116 |  | <10 |
| 110 |  | < 10 |
| 102 |  | < 5 |
| 129A |  | <0,2 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median 2,14

| Lab. no. | Dev. from median% | pg/m3 |
| --- | --- | --- |
| 113 | 55 | 3,312 |
| 131 | 17 | 2,5 |
| 115 | 12 | 2,4 |
| 131 | 7 | 2,3 |
| 103 | 7 | 2,281333333 |
| 118 | 0 | 2,14 |
| 123 | -16 | 1,8 |
| 104 | -22 | 1,666666667 |
| 105 | -25 | 1,6 |
| 129A | -25 | 1,6 |
| 109 | -32 | 1,45 |
| 108 |  | <88,884 |
| 116 |  | <10 |
| 110 |  | <10 |
| 102 |  | <5 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 14	Median: 0,74

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 123 | 265 | 2,7 |
| 129A | 76 | 1,3 |
| 109 | 62 | 1,2 |
| 103 | 33 | 0,981333333 |
| 131 | 0 | 0,74 |
| 115 | -5 | 0,7 |
| 118 | -5 | 0,7 |
| 131 | -7 | 0,69 |
| 104 | -32 | 0,5 |
| 108 |  | <262,92 |
| 110 |  | <100 |
| 116 |  | <10 |
| 102 |  | <5 |
| 105 |  | <LOD |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,92

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 129A | 216 | 2,9 |
| 115 | 20 | 1,1 |
| 109 | 17 | 1,07 |
| 103 | 11 | 1,016 |
| 118 | -11 | 0,82 |
| 131 | -14 | 0,79 |
| 131 | -15 | 0,78 |
| 104 | -67 | 0,3 |
| 108 |  | <177,768 |
| 110 |  | <100 |
| 116 |  | <10 |
| 102 |  | <5 |
| 105 |  | <LOD |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 1,3

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 115 | 7,69 | 1,4 |
| 129A | 7,69 | 1,4 |
| 118 | 4,62 | 1,36 |
| 104 | 0,00 | 1,3 |
| 103 | -14,97 | 1,105333333 |
| 131 | -23,85 | 0,99 |
| 131 | -24,62 | 0,98 |
| 108 |  | <43,82 |
| 110 |  | <100 |
| 116 |  | <10 |
| 109 |  | <5,3 (1) |
| 102 |  | <5 |
| 105 |  | <LOD |
| 101 |  | n.d |

(1) Interference

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 14	Median: 1,13

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 123 | 404 | 5,7 |
| 118 | 40 | 1,58 |
| 104 | 15 | 1,3 |
| 115 | 6 | 1,2 |
| 131 | 0 | 1,13 |
| 105 | -3 | 1,1 |
| 131 | -4 | 1,09 |
| 103 | -10 | 1,02 |
| 129A | -73 | 0,3 |
| 110 |  | <100 |
| 108 |  | <88,884 |
| 116 |  | <10 |
| 102 |  | <5 |
| 109 |  | <3,7 (1) |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 24,8

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 109 | 53 | 38 |
| 115 | 18 | 29,4 |
| 131 | 18 | 29,3 |
| 131 | 17 | 29,1 |
| 118 | 10 | 27,2 |
| 101 | 8 | 26,9 |
| 102 | 1 | 25 |
| 103 | -1 | 24,64666667 |
| 116 | -5 | 23,6 |
| 113 | -11 | 22* |
| 104 | -20 | 19,8 |
| 105 | -42 | 14,4 |
| 121 | -59 | 10,2 |
| 123 | -70 | 7,4 |
| 108 |  | <43,82 |
| 110 |  | <10 |
| 129A |  | <0,2 |

* out of calibration range

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 16,4

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 108 | 441 | 88,795116 |
| 109 | 101 | 33 |
| 116 | 69 | 27,7 |
| 131 | 12 | 18,4 |
| 131 | 11 | 18,2 |
| 115 | 8 | 17,7 |
| 101 | 2 | 16,7 |
| 118 | 0 | 16,4 |
| 102 | -2 | 16 |
| 121 | -9 | 15 |
| 103 | -16 | 13,74933333 |
| 113 | -25 | 12,336 |
| 104 | -30 | 11,5 |
| 105 | -66 | 5,6 |
| 123 | -94 | 0,92 |
| 110 |  | <10 |
| 129A |  | <0,2 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 13,25

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 116 | 71 | 22,6 |
| 109 | 62 | 21,5 |
| 102 | 51 | 20 |
| 131 | 9 | 14,5 |
| 115 | 4 | 13,8 |
| 131 | 4 | 13,8 |
| 118 | 3 | 13,6 |
| 101 | -3 | 12,9 |
| 103 | -12 | 11,72266667 |
| 105 | -23 | 10,2 |
| 113 | -24 | 10,03 |
| 104 | -26 | 9,8 |
| 123 | -65 | 4,6 |
| 121 | -83 | 2,2 |
| 108 |  | <43,82 |
| 110 |  | <10 |
| 129A |  | <0,2 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 22,5

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 116 | 192 | 65,7 |
| 108 | 115 | 48,352896 |
| 115 | 5 | 23,7 |
| 109 | 4 | 23,5 |
| 131 | 4 | 23,3 |
| 102 | 2 | 23 |
| 131 | 2 | 23 |
| 118 | 0 | 22,5 |
| 101 | -11 | 20,1 |
| 103 | -12 | 19,8 |
| 113 | -26 | 16,544 |
| 104 | -27 | 16,4 |
| 105 | -67 | 7,5 |
| 121 | -76 | 5,4 |
| 123 | -88 | 2,8 |
| 110 |  | <10 |
| 129A |  | <0,2 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 28,7

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 108 | 501 | 172,47552 |
| 129A | 112 | 60,9 |
| 104 | 73 | 49,7 |
| 109 | 67 | 48 |
| 115 | 41 | 40,4 |
| 102 | 18 | 34 |
| 101 | 10 | 31,7 |
| 116 | 1 | 28,9 |
| 118 | -1 | 28,5 |
| 103 | -3 | 27,97866667 |
| 113 | -6 | 27,1* |
| 105 | -10 | 25,7 |
| 131 | -18 | 23,5 |
| 131 | -37 | 18,1 |
| 121 | -69 | 8,8 |
| 123 | -80 | 5,8 |
| 110 |  | <10 |

*Out of calib range

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 61,9

| Lab. no. | Dev. from median | pg/m3 |
| --- | --- | --- |
| 108 | 193 | 181,145592 |
| 129A | 128 | 141,3 |
| 109 | 70 | 105 |
| 115 | 50 | 93,1 |
| 101 | 14 | 70,3 |
| 116 | 9 | 67,4 |
| 102 | 3 | 64 |
| 118 | 0 | 62,1 |
| 104 | 0 | 61,7 |
| 105 | -3 | 60,3 |
| 103 | -4 | 59,69466667 |
| 113 | -18 | 50,8 |
| 131 | -21 | 49,2 |
| 131 | -21 | 49,2 |
| 121 | -76 | 15 |
| 123 | -96 | 2,7 |
| 110 |  | <10 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 7,56

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 109 | 79 | 13,5 |
| 121 | 59 | 12 |
| 105 | 51 | 11,4 |
| 131 | 31 | 9,92 |
| 103 | 8 | 8,183321067 |
| 101 | 1 | 7,6 |
| 118 | -1 | 7,52 |
| 115 | -5 | 7,2 |
| 131 | -6 | 7,09 |
| 104 | -15 | 6,454545455 |
| 129A | -22 | 5,9 |
| 123 | -79 | 1,6 |
| 108 |  | <175,28 |
| 102 |  | <20 |
| 110 |  | <10 |
| 116 |  | <10 |
| 113 |  | interference |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories:16	Median: 7,34

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 121 | 853 | 70 |
| 123 | 134 | 17,2 |
| 105 | 59 | 11,7 |
| 109 | 57 | 11,5 |
| 101 | 8 | 7,9 |
| 103 | 1 | 7,4051152 |
| 118 | -1 | 7,28 |
| 115 | -11 | 6,5 |
| 131 | -12 | 6,44 |
| 104 | -16 | 6,166666667 |
| 131 | -17 | 6,07 |
| 129A | -25 | 5,5 |
| 108 |  | <44,442 |
| 102 |  | <20 |
| 110 |  | <10 |
| 116 |  | <10 |
| 113 |  | interference |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 6,73

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 121 | 167 | 18 |
| 105 | 106 | 13,9 |
| 129A | 49 | 10 |
| 109 | 47 | 9,9 |
| 101 | 13 | 7,6 |
| 131 | 2 | 6,87 |
| 103 | -2 | 6,596784267 |
| 131 | -6 | 6,34 |
| 104 | -10 | 6,090909091 |
| 115 | -14 | 5,8 |
| 118 | -21 | 5,3 |
| 123 | -79 | 1,4 |
| 108 |  | <43,82 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | <10 |
| 113 |  | interference |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 5,44

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 121 | 1113 | 66 |
| 105 | 152 | 13,7 |
| 129A | 47 | 8 |
| 109 | 33 | 7,25 |
| 101 | 29 | 7 |
| 103 | 7 | 5,801125333 |
| 104 | -7 | 5,083333333 |
| 131 | -15 | 4,62 |
| 131 | -16 | 4,56 |
| 115 | -17 | 4,5 |
| 118 | -22 | 4,22 |
| 123 | -61 | 2,1 |
| 108 |  | <266,652 |
| 102 |  | <10 |
| 110 |  | <10 |
| 116 |  | <10 |
| 113 |  | interference |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 4,10

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 123 | 364 | 19 |
| 121 | 144 | 10 |
| 105 | 127 | 9,3 |
| 129A | 54 | 6,3 |
| 109 | 17 | 4,8 |
| 113 | 2 | 4,199 |
| 131 | -2 | 4,0 |
| 103 | -5 | 3,9060496 |
| 131 | -5 | 3,9 |
| 118 | -6 | 3,84 |
| 115 | -17 | 3,4 |
| 104 | -51 | 2 |
| 108 |  | <43,82 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | <10 |
| 101 |  | < 6 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median: 2,768

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 121 | 623 | 20 |
| 105 | 225 | 9 |
| 109 | 61 | 4,45 |
| 129A | 1 | 2,8 |
| 103 | 0 | 2,7805264 |
| 113 | 0 | 2,768 |
| 118 | -5 | 2,64 |
| 131 | -7 | 2,58 |
| 115 | -10 | 2,5 |
| 131 | -14 | 2,38 |
| 104 | -49 | 1,416666667 |
| 108 |  | <44,442 |
| 102 |  | <10 |
| 110 |  | <10 |
| 116 |  | <10 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median: 1,69

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 113 | 76 | 2,975 |
| 105 | 66 | 2,8 |
| 121 | 36 | 2,3 |
| 129A | 24 | 2,1 |
| 103 | 21 | 2,039692 |
| 131 | 0 | 1,69 |
| 118 | -1 | 1,68 |
| 104 | -3 | 1,636363636 |
| 115 | -11 | 1,5 |
| 109 | -14 | 1,45 |
| 131 | -16 | 1,42 |
| 108 |  | <43,82 |
| 102 |  | < 20 |
| 110 |  | <10 |
| 116 |  | <10 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 1,2

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 121 | 1650 | 21 |
| 123 | 467 | 6,8 |
| 105 | 117 | 2,6 |
| 113 | 65 | 1,984 |
| 103 | 29 | 1,543156 |
| 109 | 4 | 1,25 |
| 131 | -4 | 1,15 |
| 115 | -8 | 1,1 |
| 104 | -10 | 1,083333333 |
| 118 | -13 | 1,04 |
| 131 | -18 | 0,98 |
| 129A | -58 | 0,5 |
| 108 |  | <44,442 |
| 102 |  | <20 |
| 110 |  | <10 |
| 116 |  | <10 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 16	Median: 2,24

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 113 | 380 | 10,744 |
| 105 | 168 | 6 |
| 121 | 97 | 4,4 |
| 123 | 88 | 4,2 |
| 129A | 12 | 2,5 |
| 104 | 2 | 2,272727273 |
| 109 | -2 | 2,2 |
| 131 | -7 | 2,09 |
| 118 | -8 | 2,06 |
| 131 | -9 | 2,04 |
| 103 | -17 | 1,8503672 |
| 115 | -51 | 1,1 |
| 108 |  | <43,82 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | < 5 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median: 1,23

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 121 | 1201 | 16 |
| 105 | 355 | 5,6 |
| 113 | 225 | 4 |
| 109 | 22 | 1,5 |
| 104 | 8 | 1,333333333 |
| 131 | 0 | 1,23 |
| 103 | -7 | 1,1381128 |
| 118 | -9 | 1,12 |
| 131 | -9 | 1,12 |
| 115 | -43 | 0,7 |
| 129A | -59 | 0,5 |
| 108 |  | <44,442 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | <5 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median: 2,75

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 113 | 237 | 9,265 |
| 121 | 122 | 6,1 |
| 105 | 53 | 4,2 |
| 131 | 12 | 3,08 |
| 131 | 8 | 2,98 |
| 109 | 0 | 2,75 |
| 129A | -9 | 2,5 |
| 118 | -11 | 2,46 |
| 104 | -14 | 2,363636364 |
| 103 | -20 | 2,200608 |
| 115 | -31 | 1,9 |
| 108 |  | <43,82 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | <5 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median: 1,83

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 108 | 4490 | 83,99538 |
| 121 | 938 | 19 |
| 105 | 108 | 3,8 |
| 113 | 87 | 3,424 |
| 109 | 26 | 2,3 |
| 118 | 2 | 1,86 |
| 131 | -2 | 1,8 |
| 131 | -5 | 1,73 |
| 104 | -9 | 1,666666667 |
| 129A | -23 | 1,4 |
| 103 | -31 | 1,2705984 |
| 115 | -34 | 1,2 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | <5 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median: 1,23

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 113 | 521 | 7,616 |
| 109 | 234 | 4,1 |
| 121 | 177 | 3,4 |
| 105 | 30 | 1,6 |
| 103 | 17 | 1,433796 |
| 118 | -17 | 1,02 |
| 104 | -18 | 1 |
| 131 | -26 | 0,91 |
| 131 | -33 | 0,82 |
| 115 | -35 | 0,8 |
| 108 |  | <43,82 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | <2 |
| 129A |  | <0,2 |
| 101 |  | n.d |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in pg/m3. Median given as a horizontal line.

Number of laboratories: 15	Median: 0,68

| Lab. no. | Dev. from median % | pg/m3 |
| --- | --- | --- |
| 121 | 2850 | 20 |
| 109 | 1021 | 7,6 |
| 113 | 273 | 2,528 |
| 105 | 92 | 1,3 |
| 103 | 26 | 0,856076 |
| 115 | -26 | 0,5 |
| 118 | -29 | 0,48 |
| 131 | -32 | 0,46 |
| 131 | -40 | 0,41 |
| 104 | -41 | 0,4 |
| 108 |  | <44,442 |
| 110 |  | <10 |
| 116 |  | <10 |
| 102 |  | <2 |
| 129A |  | <0,2 |
| 101 |  | n.d |

- PAH compounds, round 2,
sample 1 and 2

The following pages present the results from round 2 graphically. The results for each compound are shown on two pages with the left-hand page giving the sample 1 results and the right-hand side the corresponding data from sample 2.

The results have been given as deviations from the median in per cent (upper graph) and as concentration based on a 500 m3 sample volume (lower graph).

	HPLC

	GC/MS

GC/FID

	TLC

	HPLC and GC/MS

	GC/FID and GC/MS

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 11	Median: 9,63

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 368 | 45,04 |
| 129A | 41 | 13,6072 |
| 116 | 17 | 11,3 |
| 131 | 13 | 10,9 |
| 131 | 12 | 10,8 |
| 101 | 10 | 10,61 |
| 105 | -10 | 8,656 |
| 113 | -21 | 7,6 |
| 110 | -22 | 7,51 |
| 104 | -27 | 7,07852 |
| 123 | -44 | 5,4 |
| 109 | -57 | 4,1 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 12	Median: 0,78

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 410 | 3,978918 |
| 123 | 54 | 1,2 |
| 102 | 24 | 0,97 |
| 101 | 21 | 0,94 |
| 110 | 18 | 0,922 |
| 129A | 8 | 0,84 |
| 131 | 0 | 0,78 |
| 131 | -3 | 0,76 |
| 116 | -3 | 0,754 |
| 105 | -4 | 0,75 |
| 104 | -25 | 0,58116 |
| 109 | -35 | 0,505 |
| 113 | -44 | 0,44 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 11	Median: 3,91

| Lab. no. | Dev from median % | ng/m3 |
| --- | --- | --- |
| 108 | 190 | 11,35 |
| 131 | 28 | 5,03 |
| 131 | 28 | 5,01 |
| 101 | 8 | 4,23 |
| 123 | 7 | 4,2 |
| 129A | 2 | 3,9806 |
| 109 | -2 | 3,85 |
| 116 | -7 | 3,66 |
| 107 | -23 | 3 |
| 103 | -27 | 2,846526667 |
| 105 | -29 | 2,774 |
| 104 | -40 | 2,33984 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 12	Median: 0,281

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 249 | 0,98 |
| 123 | 188 | 0,81 |
| 109 | 80 | 0,505 |
| 131 | 21 | 0,34 |
| 131 | 14 | 0,32 |
| 129A | 1 | 0,2832 |
| 116 | 0 | 0,281 |
| 101 | 0 | 0,28 |
| 102 | -7 | 0,26 |
| 103 | -13 | 0,244963333 |
| 104 | -35 | 0,18259 |
| 107 | -36 | 0,18 |
| 105 | -43 | 0,16 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 12	Median: 1,13

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 319 | 4,74 |
| 109 | 214 | 3,55 |
| 123 | 38 | 1,56 |
| 129A | 25 | 1,4172 |
| 101 | 24 | 1,4 |
| 107 | 2 | 1,15 |
| 131 | 0 | 1,13 |
| 131 | -1 | 1,12 |
| 116 | -2 | 1,11 |
| 105 | -25 | 0,853 |
| 104 | -25 | 0,84534 |
| 113 | -32 | 0,77 |
| 103 | -40 | 0,67909 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,315

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 281 | 1,2 |
| 108 | 246 | 1,09 |
| 123 | 33 | 0,42 |
| 101 | 24 | 0,39 |
| 129A | 15 | 0,3612 |
| 116 | 10 | 0,345 |
| 131 | 2 | 0,32 |
| 131 | -2 | 0,31 |
| 107 | -24 | 0,24 |
| 104 | -25 | 0,23712 |
| 105 | -25 | 0,235 |
| 113 | -37 | 0,2 |
| 103 | -41 | 0,18725 |
| 102 | -94 | 0,02 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 12	Median: 4,31

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 778 | 37,83 |
| 129A | 101 | 8,6494 |
| 131 | 18 | 5,08 |
| 131 | 12 | 4,84 |
| 101 | 8 | 4,64 |
| 103 | 4 | 4,462316 |
| 116 | 0 | 4,31 |
| 109 | -1 | 4,25 |
| 113 | -30 | 3 |
| 105 | -32 | 2,925 |
| 107 | -36 | 2,77 |
| 104 | -46 | 2,31453 |
| 123 | -58 | 1,8 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 2,61

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 485 | 15,26 |
| 129A | 32 | 3,4548 |
| 131 | 31 | 3,43 |
| 131 | 29 | 3,36 |
| 101 | 25 | 3,25 |
| 116 | 7 | 2,78 |
| 109 | 3 | 2,7 |
| 103 | -3 | 2,517484 |
| 107 | -16 | 2,2 |
| 113 | -20 | 2,1 |
| 105 | -27 | 1,912 |
| 104 | -38 | 1,61393 |
| 102 | -48 | 1,36 |
| 123 | -75 | 0,65 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 12 	Median: 8,23

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 101 | 31 | 10,75 |
| 116 | 23 | 10,1 |
| 129A | 20 | 9,8418 |
| 131 | 4 | 8,58 |
| 109 | 3 | 8,5 |
| 131 | 3 | 8,48 |
| 107 | -3 | 7,97 |
| 105 | -17 | 6,826 |
| 103 | -18 | 6,739749333 |
| 113 | -26 | 6,1 |
| 104 | -46 | 4,42796 |
| 123 | -54 | 3,8 |
| 108 |  | < 38,65 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 6,68

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 101 | 39 | 9,28 |
| 129A | 27 | 8,4546 |
| 116 | 25 | 8,35 |
| 131 | 12 | 7,48 |
| 131 | 9 | 7,27 |
| 109 | 2 | 6,8 |
| 108 | 0 | 6,69 |
| 102 | 0 | 6,66 |
| 103 | -4 | 6,402977333 |
| 105 | -14 | 5,722 |
| 107 | -17 | 5,54 |
| 113 | -22 | 5,2 |
| 104 | -43 | 3,80953 |
| 123 | -58 | 2,8 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,80

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 613 | 5,7 |
| 113 | 338 | 3,5 |
| 101 | 60 | 1,28 |
| 129A | 31 | 1,0488 |
| 116 | 20 | 0,963 |
| 131 | 1 | 0,81 |
| 131 | 0 | 0,80 |
| 107 | -3 | 0,78 |
| 110 | -5 | 0,763 |
| 103 | -17 | 0,667454667 |
| 105 | -27 | 0,583 |
| 104 | -33 | 0,53571 |
| 123 | -69 | 0,25 |
| 108 |  | < 0,77 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 14	Median: 0,475

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 847 | 4,5 |
| 113 | 532 | 3 |
| 123 | 311 | 1,95 |
| 101 | 79 | 0,85 |
| 129A | 35 | 0,6414 |
| 116 | 26 | 0,599 |
| 131 | 1 | 0,48 |
| 131 | -1 | 0,47 |
| 110 | -4 | 0,454 |
| 103 | -6 | 0,44492 |
| 102 | -7 | 0,44 |
| 105 | -28 | 0,343 |
| 107 | -35 | 0,31 |
| 104 | -35 | 0,30811 |
| 108 |  | < 0,78 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 2,75

| Lab. no. | Dev. from median% | ng/m3 |
| --- | --- | --- |
| 110 | 99 | 5,55 |
| 116 | 20 | 3,33 |
| 101 | 10 | 3,06 |
| 129A | 7 | 2,9914 |
| 131 | 2 | 2,85 |
| 131 | 1 | 2,82 |
| 109 | -1 | 2,75 |
| 107 | -3 | 2,71 |
| 103 | -14 | 2,391342667 |
| 113 | -21 | 2,2 |
| 105 | -24 | 2,123 |
| 104 | -39 | 1,68749 |
| 123 | -68 | 0,9 |
| 108 |  | < 5,41 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 14	Median: 2,245

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 110 | 62 | 3,63 |
| 101 | 8 | 2,42 |
| 129A | 5 | 2,363 |
| 131 | 2 | 2,28 |
| 102 | 1 | 2,27 |
| 116 | 1 | 2,27 |
| 109 | 0 | 2,25 |
| 131 | 0 | 2,24 |
| 103 | -5 | 2,123284 |
| 113 | -24 | 1,7 |
| 105 | -27 | 1,649 |
| 107 | -31 | 1,56 |
| 104 | -40 | 1,34752 |
| 123 | -81 | 0,42 |
| 108 |  | < 4,68 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 2,35

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 98 | 4,65 |
| 116 | 23 | 2,9 |
| 110 | 22 | 2,86 |
| 129A | 12 | 2,6376 |
| 101 | 11 | 2,62 |
| 113 | 2 | 2,4 |
| 131 | 0 | 2,35 |
| 131 | -1 | 2,33 |
| 107 | -5 | 2,23 |
| 103 | -18 | 1,931886667 |
| 105 | -23 | 1,81 |
| 104 | -33 | 1,56447 |
| 123 | -79 | 0,5 |
| 108 |  | < 2,32 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 14	Median: 1,87

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 95 | 3,65 |
| 123 | 20 | 2,25 |
| 101 | 11 | 2,08 |
| 110 | 8 | 2,02 |
| 129A | 7 | 1,994 |
| 116 | 2 | 1,91 |
| 113 | 2 | 1,9 |
| 131 | -2 | 1,84 |
| 131 | -4 | 1,8 |
| 103 | -5 | 1,779076 |
| 105 | -27 | 1,374 |
| 107 | -31 | 1,29 |
| 104 | -36 | 1,20284 |
| 102 | -45 | 1,03 |
| 108 |  | < 2,34 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,42

| Lab. no. | Dev. from median | ng/m3 |
| --- | --- | --- |
| 108 | 243 | 1,44 |
| 116 | 60 | 0,671 |
| 129A | 11 | 0,4648 |
| 131 | 10 | 0,46 |
| 123 | 7 | 0,45 |
| 131 | 7 | 0,45 |
| 101 | 0 | 0,42 |
| 103 | -1 | 0,417665333 |
| 104 | -11 | 0,37534 |
| 113 | -17 | 0,35 |
| 110 | -19 | 0,34 |
| 105 | -23 | 0,322 |
| 107 | -40 | 0,25 |
| 109 |  | <0,21 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,27

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 263 | 0,98 |
| 116 | 17 | 0,316 |
| 103 | 15 | 0,309314667 |
| 131 | 11 | 0,3 |
| 131 | 11 | 0,3 |
| 129A | 6 | 0,2874 |
| 101 | 0 | 0,27 |
| 104 | -16 | 0,22721 |
| 105 | -21 | 0,214 |
| 110 | -26 | 0,2 |
| 113 | -26 | 0,2 |
| 107 | -33 | 0,18 |
| 102 | -89 | 0,03 |
| 109 |  | <0,21 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories:13	Median: 0,78

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 176 | 2,15 |
| 109 | 124 | 1,75 |
| 116 | 42 | 1,11 |
| 103 | 10 | 0,854905333 |
| 113 | 9 | 0,85 |
| 129A | 7 | 0,8328 |
| 131 | 1 | 0,79 |
| 131 | -1 | 0,77 |
| 101 | -17 | 0,65 |
| 107 | -32 | 0,53 |
| 105 | -36 | 0,498 |
| 110 | -43 | 0,448 |
| 104 | -48 | 0,40931 |
| 123 | -62 | 0,3 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,536

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 223 | 1,73 |
| 109 | 217 | 1,7 |
| 129A | 15 | 0,6158 |
| 131 | 14 | 0,61 |
| 131 | 12 | 0,6 |
| 113 | 10 | 0,59 |
| 116 | 1 | 0,54 |
| 103 | -1 | 0,532266667 |
| 102 | -1 | 0,53 |
| 101 | -10 | 0,48 |
| 107 | -24 | 0,41 |
| 105 | -29 | 0,379 |
| 110 | -32 | 0,362 |
| 104 | -46 | 0,29007 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 11	Median: 0,51

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 131 | 91 | 0,97 |
| 131 | 91 | 0,97 |
| 109 | 61 | 0,815 |
| 104 | 43 | 0,72282 |
| 116 | 9 | 0,555 |
| 129A | 8 | 0,5502 |
| 110 | -8 | 0,464 |
| 103 | -10 | 0,457233333 |
| 101 | -11 | 0,45 |
| 113 | -23 | 0,39 |
| 105 | -28 | 0,365 |
| 107 | -43 | 0,29 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories:12	Median: 0,389

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 131 | 95 | 0,76 |
| 131 | 93 | 0,75 |
| 109 | 80 | 0,7 |
| 104 | 37 | 0,53254 |
| 105 | 1 | 0,391 |
| 101 | 0 | 0,39 |
| 116 | 0 | 0,389 |
| 129A | -1 | 0,386 |
| 103 | -8 | 0,35818 |
| 113 | -20 | 0,31 |
| 102 | -25 | 0,29 |
| 110 | -28 | 0,279 |
| 107 | -49 | 0,2 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 11	Median: 0,226

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 410 | 1,15 |
| 108 | 210 | 0,7 |
| 129A | 135 | 0,5302 |
| 116 | 57 | 0,355 |
| 101 | 2 | 0,23 |
| 104 | 0 | 0,22559 |
| 103 | -15 | 0,190993333 |
| 110 | -18 | 0,185 |
| 105 | -24 | 0,172 |
| 113 | -29 | 0,16 |
| 107 | -34 | 0,15 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 12	Median: 0,177

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 551 | 1,15 |
| 108 | 273 | 0,66 |
| 129A | 136 | 0,4166 |
| 116 | 55 | 0,274 |
| 101 | 2 | 0,18 |
| 103 | 1 | 0,178508 |
| 104 | -1 | 0,17506 |
| 105 | -5 | 0,168 |
| 102 | -15 | 0,15 |
| 110 | -20 | 0,142 |
| 107 | -21 | 0,14 |
| 113 | -21 | 0,14 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,36

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 276 | 1,37 |
| 122 | 95 | 0,71 |
| 109 | 43 | 0,52 |
| 116 | 27 | 0,462 |
| 129A | 13 | 0,4114 |
| 101 | 4 | 0,38 |
| 113 | 4 | 0,38 |
| 110 | -4 | 0,349 |
| 103 | -12 | 0,3223 |
| 131 | -12 | 0,32 |
| 131 | -12 | 0,32 |
| 107 | -23 | 0,28 |
| 104 | -23 | 0,2793 |
| 105 | -23 | 0,279 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 14	Median: 0,24

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 296 | 0,95 |
| 109 | 42 | 0,34 |
| 122 | 38 | 0,33 |
| 116 | 9 | 0,262 |
| 101 | 8 | 0,26 |
| 103 | 2 | 0,245026667 |
| 129A | 1 | 0,2412 |
| 102 | 0 | 0,24 |
| 110 | -16 | 0,202 |
| 131 | -17 | 0,2 |
| 131 | -17 | 0,2 |
| 113 | -21 | 0,19 |
| 105 | -22 | 0,187 |
| 104 | -29 | 0,17022 |
| 107 | -38 | 0,15 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,402

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 437 | 2,16 |
| 123 | 397 | 2 |
| 116 | 26 | 0,508 |
| 103 | 24 | 0,498177333 |
| 131 | 2 | 0,41 |
| 105 | 2 | 0,409 |
| 129A | 1 | 0,4042 |
| 131 | -1 | 0,40 |
| 110 | -5 | 0,382 |
| 109 | -8 | 0,37 |
| 101 | -13 | 0,35 |
| 113 | -18 | 0,33 |
| 104 | -22 | 0,31502 |
| 107 | -30 | 0,28 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,305

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 490 | 1,8 |
| 103 | 57 | 0,4795 |
| 116 | 28 | 0,391 |
| 109 | 20 | 0,365 |
| 105 | 18 | 0,361 |
| 102 | 11 | 0,34 |
| 131 | 2 | 0,31 |
| 131 | -2 | 0,3 |
| 110 | -2 | 0,299 |
| 129A | -8 | 0,2814 |
| 113 | -15 | 0,26 |
| 101 | -18 | 0,25 |
| 104 | -24 | 0,23082 |
| 107 | -38 | 0,19 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,05

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 123 | 3700 | 1,9 |
| 110 | 476 | 0,288 |
| 116 | 54 | 0,077 |
| 129A | 46 | 0,0732 |
| 101 | 0 | 0,05 |
| 113 | 0 | 0,05 |
| 131 | 0 | 0,05 |
| 131 | 0 | 0,05 |
| 103 | -4 | 0,048069333 |
| 105 | -6 | 0,047 |
| 104 | -22 | 0,03919 |
| 108 |  | < 6,18 |
| 109 |  | <0,26 |
| 107 |  | <LQ |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,0416

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 17737 | 7,42 |
| 109 | 657 | 0,315 |
| 102 | 68 | 0,07 |
| 110 | 20 | 0,05 |
| 103 | 9 | 0,045304 |
| 105 | 1 | 0,042 |
| 129A | 0 | 0,0416 |
| 101 | -4 | 0,04 |
| 131 | -4 | 0,04 |
| 116 | -11 | 0,037 |
| 131 | -28 | 0,03 |
| 104 | -37 | 0,02639 |
| 113 | -52 | 0,02 |
| 107 |  | <LQ |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 13	Median: 0,46

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 274 | 1,72 |
| 123 | 74 | 0,8 |
| 109 | 67 | 0,77 |
| 116 | 43 | 0,66 |
| 129A | 20 | 0,5504 |
| 105 | 2 | 0,47 |
| 113 | 0 | 0,46 |
| 131 | 0 | 0,46 |
| 101 | -2 | 0,45 |
| 131 | -2 | 0,45 |
| 103 | -6 | 0,434190667 |
| 107 | -11 | 0,41 |
| 110 | -25 | 0,347 |
| 104 | -33 | 0,30642 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 14	Median: 0,355

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 108 | 311 | 1,46 |
| 109 | 83 | 0,65 |
| 102 | 55 | 0,55 |
| 105 | 15 | 0,41 |
| 129A | 14 | 0,4054 |
| 113 | 1 | 0,36 |
| 131 | 1 | 0,36 |
| 103 | 0 | 0,355498667 |
| 116 | -1 | 0,353 |
| 131 | -2 | 0,35 |
| 101 | -4 | 0,34 |
| 107 | -21 | 0,28 |
| 110 | -22 | 0,277 |
| 104 | -37 | 0,22288 |
| 123 | -77 | 0,08 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 6	Median: 4,62

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 75 | 8,1 |
| 116 | 2 | 4,7 |
| 131 | 1 | 4,68 |
| 131 | 0 | 4,62 |
| 129A | -26 | 3,3996 |
| 104 | -50 | 2,32367 |
| 123 | -97 | 0,15 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 5	Median: 1,16

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 73 | 2 |
| 129A | 31 | 1,5104 |
| 131 | 2 | 1,18 |
| 131 | -2 | 1,13 |
| 116 | -3 | 1,12 |
| 104 | -48 | 0,60386 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 6	Median: 0,035

| Lab. no. | Dev. from median | ng/m3 |
| --- | --- | --- |
| 129A | 2489 | 0,8932 |
| 109 | 1972 | 0,715 |
| 116 | 42 | 0,049 |
| 131 | -42 | 0,02 |
| 131 | -42 | 0,02 |
| 104 | -47 | 0,01843 |
| 107 |  | <LQ |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 6	Median: 0,02

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 4150 | 0,85 |
| 116 | 90 | 0,038 |
| 131 | 0 | 0,02 |
| 131 | 0 | 0,02 |
| 104 | -6 | 0,01881 |
| 129A | -8 | 0,0184 |
| 107 |  | <LQ |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 9	Median: 0,4086

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 152 | 1,03 |
| 103 | 81 | 0,739625333 |
| 116 | 43 | 0,586 |
| 104 | 24 | 0,50804 |
| 129A | 0 | 0,4086 |
| 131 | -17 | 0,34 |
| 101 | -19 | 0,33 |
| 131 | -19 | 0,33 |
| 107 | -29 | 0,29 |
| 108 |  | < 1,55 |

Upper figure:	Deviation from median in %.

Lower figure:	Concentrations in ng/m3. Median given as a horizontal line.

Number of laboratories: 9	Median: 0,317

| Lab. no. | Dev. from median % | ng/m3 |
| --- | --- | --- |
| 109 | 263 | 1,15 |
| 103 | 79 | 0,568157333 |
| 116 | 30 | 0,411 |
| 104 | 23 | 0,38935 |
| 129A | 0 | 0,3166 |
| 131 | -15 | 0,27 |
| 101 | -18 | 0,26 |
| 131 | -18 | 0,26 |
| 107 | -31 | 0,22 |
| 108 |  | < 01,56 |