## Slide 1: Results of the 1st comparison exercise for TC, OC and EC measurements within ACTRIS + EMEP laboratories
Fabrizia Cavalli and J.-P Putaud

## Slide 2: 1. Samples, sub-samples and homogeneity
Six PM10 aerosol samples:
- SPA-1 and SPA-2: Montseny, Spain;
- HUN-1 and HUN-2: K-Puzsta, Hungary;
- ITA-1 and ITA-2: Ispra, Italy.
Prescribed scheme for punch distribution
Homogeneity test (rsd% of TC measurements) on separate filters:
- SPA-1 and SPA-2, 5%;
- HUN-1 and HUN-2, 1%;
- ITA-1 and ITA-2 2%;
If sampling at each location occurred under repeatability conditions, then comparison samples should have similar homogeneity

## Slide 3
2. Participants
2 participants used NIOSH-like protocol with 840 °C for the highest T step in the He-mode
12 participants use the EUSAAR_2 protocol with 650 °C for the highest T step in the He-mode
1 participant uses VDI method without correction for pyrolysis
2 participants determined TC only
 | Name of laboratory | ACTRIS status | Contact persons
1 | ISAC-CNR | P | A.Marinoni@isac.cnr.it
2a | LGGE grenoble | p | cozic@lgge.obs.ujf-grenoble.fr
2b | LSCE | P | jean.sciare@lsce.ipsl.fr
3 | Nilu | P | karl.yttri@nilu.no
6 | Unibe | P | soenke.szidat@iac.unibe.ch
7 | IFT | P | spindler@tropos.de
8 | Empa | P | Christoph.Hueglin@empa.ch
12 | Ulund | P | johan.genberg@nuclear.lu.se
13 | CSIC | P | mariacruz.minguillon@idaea.csic.es
16 | UOC | P | ctheodosi@gmail.com
17 | JRC | P | fabrizia.cavalli@jrc.ec.europa.eu
25 | CZE | P | novakj@chmi.cz
26 | UPAC | P | kissgy@almos.uni-pannon.hu
A10 | NCSR D | A | ldiapouli@ipta.demokritos.gr
A18 | Bham | A | R.M.HARRISON@bham.ac.uk
E1 | ISCIII | E | rferndez@isciii.es
E2 | CNR-IIA | E | perrino@iia.cnr.it
E3 | UBA-DE | E | elke.bieber@uba.de

## Slide 4
3. Data evaluation: methodology
In absence of a general consensus within the scientific community
(Baumgardener et al. Soot Reference Materials for Instrument Calibration and Intercomparisons A Workshop Summary with Recommendations, AMTD, 2012) :
- definition of EC
standard reference material
standard reference analytical method
in this interlaboratory comparison TC and EC/TC are evaluated for:
3.1 METHOD PERFORMANCE (ISO 5725-2)
▪ Mandel’s h and k statistics;
▪ Reproducibility and Repeatability standard deviations and their dependence upon the mean values.
3.2 LABORATORY PERFORMANCE (ISO 13528:2005)
▪ Assigned value: “Consensus value from participants”, i.e. robust average;
▪ Standard deviation for proficiency assessment:
  - from data obtained in a round of a proficiency testing scheme, i.e. robust standard deviation;
  - prescribed value: DQO 25% (expanded uncertainty, k=2) for PM at its limit value (EU-dir.);
  - by perception: level of performance wished to be achieved by laboratories, i.e. 15%.
▪ z-scores
EC/TC
TC

## Slide 5
On average,
TC amount range 8.7 to 37.0 μg cm-2 equivalent to 5.0 to 21.3 μg m-3;
(collected for 24h with a face velocity of 20.1 cm s-1)
EC/TC ranged from 0.02 to 0.14
No technical reasons for which different thermal protocol could result in different TC values:
discrepancies in TC considered as random and TC database is evaluated as a whole.
By contrary, OC/EC split is protocol dependent: it is more appropriate to evaluate separately two subsets according to the protocol.
However, EC/TC is evaluated as a whole as a subset of two labs is not statistically significant.

## Slide 6
3.1 METHOD PERFORMANCE (ISO 5725-2)
Mandel’s h statistics:
(x-X)/σ for each lab and within each sample
where
x = reported value; X= grand mean; and σ = standard deviation of the mean
TC – between laboratory consistency
One outlier and seven stragglers
Chart: Laboratory
| 1 | 2a | 2b | 3 | 6 | 7 | 8 | 12 | 13 | 16 | 17 | 25 | 26 | A10 | A 18 | E1 | E2 | E3
SPA-1 | 0.80516 | 0.6524 | 0.46178 | -0.14812 | 2.9233 | -1.1997 | 0.053071 | 0.54981 | 0.16441 | -0.54792 | -0.21866 | -0.3323 | -0.89393 | 1.9164 | -0.43076 | -0.11767 | 0.4491 | -0.20687
SPA-2 | 1.395 | 0.78832 | 0.0052493 | -0.21494 | -0.52399 | -1.9189 | 0.54202 | 1.3769 | 0.33188 | -0.98748 | 0.13484 | -0.5088 | 0.31567 | 2.0161 | 0.055659 | -1.0405 | 1.2556 | -0.22541
HUN-1 | 0.56535 | 0.71584 | -0.12429 | -0.074923 | -0.11634 | -2.1047 | 1.0786 | 0.74765 | 0.3352 | -0.95469 | 0.21993 | -0.71986 | -0.47666 | 1.8028 | -0.099024 | 0.43528 | 2.1689 | -0.10067
HUN-2 | 0.23477 | 0.64958 | -0.14342 | 0.70744 | -0.10425 | -2.0086 | 0.20558 | 0.88027 | 1.0986 | -1.1832 | -0.4119 | -0.33158 | -0.096792 | 1.9382 | -0.23795 | 0.52431 | 2.0366 | 0.087082
ITA-1 | -0.37292 | -0.71037 | -0.70149 | 0.22612 | 1.4023 | -0.8812 | 0.98975 | 1.3094 | -0.14368 | -0.90381 | -0.38575 | -0.92503 | -0.83389 | 1.23 | -0.89763 | -0.063395 | 2.1542 | 0.93554
ITA-2 | 0.90171 | 1.3471 | 1.1434 | 1.1974 | 1.0632 | -1.7289 | -0.33747 | -0.35885 | -0.38195 | -1.0048 | -0.7883 | -1.0122 | -0.30794 | 0.22455 | 1.0523 | 1.241 | 1.189 | -0.51097
 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36 | -2.36
 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88 | -1.88
 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88 | 1.88
 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36 | 2.36

## Slide 7
TC – within laboratory consistency
3.1 METHOD PERFORMANCE (ISO 5725-2)
Six outliers and one strugglers
Chart: Laboratory
| 1 | 2a | 2b | 3 | 6 | 7 | 8 | 12 | 13 | 16 | 17 | 25 | 26 | A10 | A 18 | E1 | E2 | E3
SPA-1 | 1.0677 | 0.58837 | 0 | 0.67813 | 1.6547 | 1.595 | 0.37885 | 0 | 0 | 0.3176 | 0.55853 | 0.68496 | 2.4779 | 0 | 0.38942 | 0.064798 | 0.60911 | 0.28148
SPA-2 | 0.64481 | 0.14457 | 0 | 0.26285 | 0.37754 | 0.76877 | 0.10052 | 0.052395 | 0 | 0.56797 | 0.14271 | 0.28761 | 1.335 | 0 | 0.13599 | 3.3599 | 0.22923 | 0.42449
HUN-1 | 0.081664 | 0.037754 | 0 | 0.61003 | 0.51011 | 1.8732 | 2.5547 | 0 | 0 | 0.078962 | 0.97362 | 0.53363 | 0.41673 | 0 | 0.81149 | 0 | 1.1141 | 0.11049
HUN-2 | 1.0803 | 0.026236 | 0 | 0.71576 | 0.66755 | 0.85887 | 0.2732 | 0 | 0 | 0.23566 | 0.50562 | 0.84606 | 3.1257 | 0 | 0.15168 | 0.47946 | 0.089499 | 0.065027
ITA-1 | 0.31258 | 0.0024306 | 0 | 0.78769 | 1.103 | 0.79186 | 0.44613 | 0 | 0 | 0.020874 | 1.405 | 0.23046 | 1.7065 | 0 | 0.45289 | 2.3385 | 0.7284 | 0.30814
ITA-2 | 2.7189 | 0.0051651 | 0 | 0.33895 | 0.40758 | 0.66071 | 0.15903 | 0.14115 | 0 | 0.2652 | 0.15137 | 0.075174 | 0.73329 | 0 | 0.11932 | 0.20827 | 2.4725 | 0.19361
 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41 | 2.41
 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93 | 1.93
Mandel’s k statistics:
si ▪ √p / √Σsi for each lab, i, within each sample
where
s = reported standard deviation; and p = number of participant

## Slide 8
A dependence of precision upon the mean values (weighted linear relationship):
sr = 0.024 × m + 0.183;
sR = 0.169 × m - 0.174.
TC
Mean value, the repeatability, sr, and reproducibility, sR
3.1 METHOD PERFORMANCE (ISO 5725-2)
Chart: m
| 8.1377035652009297 | 14.951326207141902 | 17.507139293557351 | 20.54634829502864 | 36.374293531152546 | 29.499043461053759
sR | 1.435 | 2.0799 | 2.6441 | 2.5699 | 8.256 | 8.0727
sr | 0.44398 | 0.54054 | 0.51948 | 0.64867 | 1.8914 | 0.83604
Worse precision values particularly for ITA-1 and ITA-2 might be due to filter heterogeneity
 | General mean (µg/cm²) | sr (%) | sR (%)
SPA-1 | 8.14 | 5.5 | 17.6
SPA-2 | 14.95 | 3.6 | 13.9
HUN-1 | 17.51 | 3.0 | 15.1
HUN-2 | 20.55 | 3.2 | 12.5
ITA-1 | 36.37 | 5.2 | 22.7
ITA-2 | 29.50 | 2.8 | 27.4

## Slide 9
Chart: sample
| SPA-1 | SPA-2 | HUN-1 | HUN-2 | ITA-1 | ITA-2
1 | 0.80516 | 1.395 | 0.56535 | 0.23477 | -0.37292 | 0.90171
2a | 0.6524 | 0.78832 | 0.71584 | 0.64958 | -0.71037 | 1.3471
2b | 0.46178 | 0.0052493 | -0.12429 | -0.14342 | -0.70149 | 1.1434
3 | -0.14812 | -0.21494 | -0.074923 | 0.70744 | 0.22612 | 1.1974
6 | 2.9233 | -0.52399 | -0.11634 | -0.10425 | 1.4023 | 1.0632
7 | -1.1997 | -1.9189 | -2.1047 | -2.0086 | -0.8812 | -1.7289
8 | 0.053071 | 0.54202 | 1.0786 | 0.20558 | 0.98975 | -0.33747
12 | 0.54981 | 1.3769 | 0.74765 | 0.88027 | 1.3094 | -0.35885
13 | 0.16441 | 0.33188 | 0.3352 | 1.0986 | -0.14368 | -0.38195
16 | -0.54792 | -0.98748 | -0.95469 | -1.1832 | -0.90381 | -1.0048
17 | -0.21866 | 0.13484 | 0.21993 | -0.4119 | -0.38575 | -0.7883
25 | -0.3323 | -0.5088 | -0.71986 | -0.33158 | -0.92503 | -1.0122
Chart: sample
| SPA-1 | SPA-2 | HUN-1 | HUN-2 | ITA-1 | ITA-2
1 | 1.0677 | 0.64481 | 0.081664 | 1.0803 | 0.31258 | 2.7189
2a | 0.58837 | 0.14457 | 0.037754 | 0.026236 | 0.0024306 | 0.0051651
2b | 0 | 0 | 0 | 0 | 0 | 0
3 | 0.67813 | 0.26285 | 0.61003 | 0.71576 | 0.78769 | 0.33895
6 | 1.6547 | 0.37754 | 0.51011 | 0.66755 | 1.103 | 0.40758
7 | 1.595 | 0.76877 | 1.8732 | 0.85887 | 0.79186 | 0.66071
8 | 0.37885 | 0.10052 | 2.5547 | 0.2732 | 0.44613 | 0.15903
12 | 0 | 0.052395 | 0 | 0 | 0 | 0.14115
13 | 0 | 0 | 0 | 0 | 0 | 0
16 | 0.3176 | 0.56797 | 0.078962 | 0.23566 | 0.020874 | 0.2652
17 | 0.55853 | 0.14271 | 0.97362 | 0.50562 | 1.405 | 0.15137
25 | 0.68496 | 0.28761 | 0.53363 | 0.84606 | 0.23046 | 0.075174
TC – within laboratory consistency
TC – between laboratory consistency
The recurrence of straggler/outliers for a single lab (for other samples beside ITA-1 and ITA-2) indicate an unsatisfactory lab reproducibility or repeatability as compared to other labs.

## Slide 10
3.1 METHOD PERFORMANCE (ISO 5725-2)
Mandel’s h statistics on the entire dataset
EC/TC – between laboratory consistency
Lab 7 (VDI method without pyrolysis correction - EC/TC higher by up to a factor 6) excluded from evaluation
Labs A10 and E1 used a NISOH-like protocol, the remaining labs the EUSAAR_2 protocol
One outlier and one straggler
Chart: Laboratory
| 1 | 2a | 2b | 3 | 6 | 7 | 8 | 12 | 13 | 16 | 17 | 25 | 26 | A10 | A 18 | E1 | E2 | E3
SPA-1 | 0 | 1.5047 | -1.1794 | -0.65085 | 0.80676 | 0 | 0.84604 | -1.1873 | -1.5945 | 0.61151 | -0.70913 | 0.25761 | 0 | 1.2829 | 0.44785 | -0.31985 | -1.3072 | 0.04029
SPA-2 | 0 | 1.8541 | -1.1132 | 0.40513 | 0.5631 | 0 | 0.52476 | -1.692 | 0.17448 | 0.067203 | -0.51679 | -0.12769 | 0 | -1.2915 | 0.83812 | 0.42952 | -1.6227 | 0.44416
HUN-1 | 0 | 1.0936 | 0.81265 | -1.0919 | -1.1153 | 0 | -0.81787 | -1.697 | 0.81458 | 0.36919 | 1.3638 | -0.53764 | 0 | 0.123 | 0.44722 | 1.5159 | -0.76046 | -0.34453
HUN-2 | 0 | 1.7917 | 0.262 | -1.1996 | -1.5032 | 0 | 0.017549 | -1.1179 | 1.248 | 0.19645 | 0.89773 | -1.1768 | 0 | 0.18406 | 1.1865 | 0.025914 | -0.72885 | -0.11099
ITA-1 | 0 | 0.23721 | -0.95869 | -0.79931 | 1.4342 | 0 | 0.13572 | -0.97002 | -1.3757 | 1.0962 | -0.32545 | -0.13081 | 0 | -0.81465 | -0.27196 | -0.67709 | 2.2109 | -0.26064
ITA-2 | 0 | -0.11346 | -0.62674 | -0.64346 | -0.88026 | 0 | 0.22535 | -0.12237 | -0.67804 | 0.36151 | -0.076183 | -0.64311 | 0 | 0.14198 | -0.58209 | 0.013919 | 3.312 | -0.036687
 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32 | -2.32
 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86 | -1.86
 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32 | 2.32
 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86 | 1.86

## Slide 11
3.1 METHOD PERFORMANCE (ISO 5725-2)
Mandel’s k statistics
EC/TC – within laboratory consistency
Chart: Laboratory
| 1 | 2a | 2b | 3 | 6 | 7 | 8 | 12 | 13 | 16 | 17 | 25 | 26 | A10 | A 18 | E1 | E2 | E3
SPA-1 | 0 | 0.79853 | 0 | 0.34001 | 1.8361 | 0 | 0.34816 | 0 | 0 | 0.53764 | 0.5216 | 1.8505 | 0 | 0 | 0.34159 | 1.402 | 0.28706 | 0.77727
SPA-2 | 0 | 0.22425 | 0 | 0.61318 | 0.1708 | 0 | 0.23007 | 0.76746 | 0 | 0.39739 | 0.27461 | 0.27843 | 0 | 0 | 0.029772 | 2.3883 | 0.64965 | 2.113
HUN-1 | 0 | 0.30735 | 0 | 0.30413 | 0.74963 | 0 | 1.7488 | 0 | 0 | 0.97281 | 0.44837 | 0.8131 | 0 | 0 | 0.089329 | 1.1775 | 0.87385 | 1.7961
HUN-2 | 0 | 1.3656 | 0 | 0.16688 | 1.0521 | 0 | 0.73176 | 0 | 0 | 0.93254 | 0.73309 | 1.9156 | 0 | 0 | 0.39765 | 0.5096 | 0.15901 | 1.3947
ITA-1 | 0 | 0.19923 | 0 | 0.15883 | 2.8382 | 0 | 0.31825 | 0 | 0 | 0.030772 | 0.51597 | 0.56047 | 0 | 0 | 0.42439 | 0.20784 | 0.097425 | 1.4015
ITA-2 | 0 | 0.071699 | 0 | 0.86582 | 0.66144 | 0 | 0.38099 | 0.14456 | 0 | 0.010002 | 0.36391 | 0.80346 | 0 | 0 | 0.1422 | 0.58628 | 2.5 | 1.8027
 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34 | 2.34
 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91 | 1.91
Three outliers and two stragglers

## Slide 12
EC/TC
Mean value, the repeatability, sr, and reproducibility, sR
 | General mean | sr (%) | sR (%)
SPA-1 | 0.11 | 7.9% | 16.5%
SPA-2 | 0.14 | 4.4% | 15.0%
HUN-1 | 0.10 | 5.7% | 11.7%
HUN-2 | 0.09 | 5.7% | 11.5%
ITA-1 | 0.04 | 11.9% | 27.6%
ITA-2 | 0.01 | 15.9% | 32.9%
No relationship between precision values and mean values
Worse precision values particularly for ITA-1 and ITA-2 due to their very low EC/TC ratio values or filter heterogeneity
The recurrence of stragglers/outliers for a single lab (for other samples beside ITA-1 and ITA-2) indicate an unsatisfactory lab reproducibility or repeatability as compared to other labs.

## Slide 13
TC
One σ represents a % deviation from the assigned value of ca. 12%;
Seventeen strugglers (of which eleven at ITA-1 + ITA-2 samples) and five outliers (of which two at ITA-1 + ITA-2 samples).
3.2 LABORATORY PERFORMANCE (ISO 13528: 2005)
z-score: X is robust average and σ is the DQO for PM of 25% at 50 µg m-3
Chart: laboratory
| 1 | 2a | 2b | 3 | 6 | 7 | 8 | 12 | 13 | 16 | 17 | 25 | 26 | A10 | A 18 | E1 | E2 | E3
SPA-1 | 1.8378 | 1.4525 | 0.97171 | -0.56652 | 7.1799 | -3.2187 | -0.059098 | 1.1937 | 0.22171 | -1.5749 | -0.74442 | -1.031 | -2.4476 | 4.6403 | -1.2794 | -0.48973 | 0.93973 | -0.71471
SPA-2 | 1.4847 | 0.75503 | -0.18683 | -0.45167 | -0.8234 | -2.5012 | 0.45878 | 1.463 | 0.20604 | -1.3809 | -0.030958 | -0.80512 | 0.18654 | 2.2318 | -0.1262 | -1.4447 | 1.3171 | -0.46427
HUN-1 | 0.50776 | 0.69591 | -0.35448 | -0.29275 | -0.34453 | -2.8305 | 1.1494 | 0.73568 | 0.22001 | -1.3927 | 0.07589 | -1.0991 | -0.79502 | 2.0549 | -0.32288 | 0.34514 | 2.5127 | -0.32494
HUN-2 | 0.032483 | 0.45151 | -0.34955 | 0.50995 | -0.30998 | -2.2337 | 0.0029976 | 0.68455 | 0.90507 | -1.3999 | -0.62076 | -0.53963 | -0.30245 | 1.7532 | -0.44504 | 0.32497 | 1.8527 | -0.11671
ITA-1 | -0.72075 | -1.3813 | -1.364 | 0.45191 | 2.7545 | -1.7158 | 1.9468 | 2.5725 | -0.27199 | -1.76 | -0.74586 | -1.8016 | -1.6231 | 2.4171 | -1.7479 | -0.11483 | 4.2263 | 1.8407
ITA-2 | 1.508 | 2.4363 | 2.0117 | 2.1242 | 1.8447 | -3.9743 | -1.0745 | -1.1191 | -1.1672 | -2.4652 | -2.0141 | -2.4808 | -1.013 | 0.09677 | 1.8219 | 2.2151 | 2.1068 | -1.4361
 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
(x-X)/σ for each lab and within each sample
where
x = reported value; X= robust average; and σ = standard deviation for proficiency assessment

## Slide 14
on two separated datasets from NIOSH-like and EUSAAR_2 users
One σ represents a % deviation from the assigned value of 15%;
Four outliers and nine strugglers - all at ITA-1+ITA-2 samples;
3.2 LABORATORY PERFORMANCE (ISO 13528: 2005)
z-score: X is robust average and σ defined by perception, i.e. 15%
EC/TC
Chart: z-scores
| 1 | 2a | 2b | 3 | 6 | 7 | 8 | 12 | 13 | 16 | 17 | 25 | 26 | A10 | A 18 | E1 | E2 | E3
SPA-1 | 0 | 1.8111 | -1.2629 | -0.65753 | 1.0118 | 0 | 1.0568 | -1.2719 | -1.7383 | 0.78818 | -0.72428 | 0.38288 | 0 | 1.5571 | 0.60075 | -0.27845 | -1.4093 | 0.13399
SPA-2 | 0 | 1.7511 | -1.1647 | 0.32729 | 0.48252 | 0 | 0.44485 | -1.7335 | 0.10064 | -0.0047726 | -0.57864 | -0.19629 | 0 | -1.3399 | 0.75277 | 0.35126 | -1.6653 | 0.36565
HUN-1 | 0 | 0.7615 | 0.56378 | -0.77679 | -0.79323 | 0 | -0.5839 | -1.2027 | 0.56513 | 0.25163 | 0.95175 | -0.38666 | 0 | 0.07835 | 0.30656 | 1.0588 | -0.5435 | -0.25073
HUN-2 | 0 | 1.3013 | 0.19143 | -0.86906 | -1.0894 | 0 | 0.01406 | -0.80983 | 0.90688 | 0.14387 | 0.65271 | -0.85253 | 0 | 0.13488 | 0.86224 | 0.020129 | -0.52752 | -0.079209
ITA-1 | 0 | 0.99177 | -1.526 | -1.1904 | 3.5117 | 0 | 0.7781 | -1.5498 | -2.404 | 2.8002 | -0.19282 | 0.21696 | 0 | -1.2227 | -0.080189 | -0.93313 | 5.1471 | -0.05636
ITA-2 | 0 | 0.35381 | -2.2941 | -2.3803 | -3.6019 | 0 | 2.1016 | 0.30784 | -2.5587 | 2.804 | 0.5461 | -2.3785 | 0 | 1.6715 | -2.0638 | 1.0109 | 18.0248 | 0.74985
 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2

## Slide 15
Prescribed scheme for punches
Method performance statistics:
TC: one outlier for reproducibility and six outliers for repeatability;
EC/TC: one outlier for reproducibility and three outliers for repeatability in the entire database;
as major filter heterogeneities could be excluded (except for ITA-1 and ITA-2), these inconsistencies indicate a poorer reproducibility/repeatability wrt. other laboratories.
TC: reproducibilty of 13-27% and repeatibility of 3-6%;
EC/TC: reproducibilty of 12-33% and repeatibility of 4-16%
Worse precision values were obtained for samples ITA-1 and ITA-2, possibly due to a poor filter homogeneity or a low value for EC/TC ratio.
Laboratory performance, z-scores:
TC: Five outliers would not comply with the DQO of 25%;
EC/TC: Four outliers do not complying with the level of performance of 15
Conclusions and recommendations
Laboratories showing biases shall carefully examine their procedures (particularly, determination and verification of the calibration constant, measurement of filter samples etc.) and identify appropriate corrective actions that are likely to prevent the recurrence of such results.