## Slide 1: Results of the inter-laboratory exercise for TC, OC and EC measurement OCEC-2016-1
Fabrizia Cavalli and Jean-Philippe Putaud

## Slide 2: Introduction
Requirements for measuring TC and EC in PM2.5 At rural background location in Air Quality Directive 2008/50/EC
Thermal-optical analysis is the most suitable method for the determination of EC and OC collected on filters and the thermal protocol EUSAAR-2 with a transmittance optical correction for pyrolysis has been selected as the European standard thermal protocol (EN16909)
The European center for aerosol calibration within the European project ACTRIS-2 has organized in 2016 an inter-laboratory comparison exercise (ILCE) (ref. OCEC-2016-1)
Notes: Total carbon (TC), including Organic Carbon (OC) and Elemental Carbon (EC) is a relevant constituent of the fine fraction of particulate matter (PM), both from the perspective of health risks due to inhalation and indication of air pollution sources. For these reasons requirements for measuring EC and OC in PM2.5 at rural background locations have been included in Air Quality Directive 2008/50/EC.
The directive states that measurements should be made in a manner consistent with those of the cooperative programme for monitoring and evaluation of the long range transmission of air pollutants in Europe (EMEP). Thermal-optical analysis has been recognized as the most suitable method for the determination of EC and OC collected on filters (see Technical Specification by the CEN/TC265 WG35) and the thermal protocol EUSAAR-2 with a transmittance optical correction for pyrolysis -already adopted in the EMEP manual for sampling and analysis- has been recently selected as the European standard thermal protocol (EN16909).
The European center for aerosol calibration within the European project ACTRIS-2 has organized in 2016 an inter-laboratory comparison exercise (ILCE) (ref. OCEC-2016-1)

## Slide 3: Organization
Ambient PM aerosol collected on quartz fiber filters at 4 sites across Europe (D, G, E, N)
The filter sample homogeneity investigated by the JRC on distinct samples collected at each location: better than 6% for TC and 11% for EC/TC
Notes: In lack of certified reference material for atmospheric OC and EC, this ILCE made use of ambient (outdoor) PM aerosol collected with high-volume samplers on quartz fiber filters at 4 sites across Europe.
The homogeneity of these test samples was investigated by the JRC on distinct samples collected at each location. From each sample, ten subsamples of 1 cm² were taken along two perpendicular axes across the filter surface and analysed for their TC, OC and EC contents. The filter homogeneity was assessed as the standard deviation of the average of the 10 replicate analyses. This leads to an upper limit for the filter homogeneity since it includes the repeatability of the JRC laboratory (< 3 and 6% for TC and EC, respectively). The homogeneity is better than 6 and 11% for TC and EC/TC, respectively (Table 2). If sampling at each location occurred under repeatable conditions, it can be assumed that the test samples had similar homogeneities.
An aqueous solution of phthalic acid was also distributed to the participants to assess the uncertainty of the instrument calibration constant determination. The solution was prepared by dissolving a precisely known mass of pure phthalic acid (≥ 99.5%) in a precisely known volume of ultra-pure water (resistivity ≥18.2 mΩ cm).

## Slide 4: Organization
Thirteen laboratories in charge of OC and EC measurements at the EMEP stations in Cyprus, Czech Republic, Finland, France, Germany, Greece, Italy, Norway, Poland, Spain and Sweden
Code | Participant | Instrument | Protocol | Punch size (cm²) | Charring cor.
1 | ACES | Sunset | EUSAAR_2 | 1.50 | Transmittance
2 | NILU | Sunset | EUSAAR_2 | 1.50 | Transmittance
3 | IEE | Sunset | EUSAAR_2 | 1.50 | Transmittance
4 | NOA | Sunset | EUSAAR_2 | n.a. | Transmittance
5 | CHMI | Sunset | EUSAAR_2 | 1.50 | Transmittance
6 | TROPOS | Sunset | EUSAAR_2 | 1.50 | Transmittance
7 | IDAEA | Sunset | EUSAAR_2 | 1.50 | Transmittance
8 | LGGE | Sunset | EUSAAR_2 | 1.50 | Transmittance
9 | UBA | Sunset | EUSAAR_2 | 1.50 | Transmittance
10 | JRC | Sunset | EUSAAR_2 | 1.00 | Transmittance
11 | Cyl | Sunset | EUSAAR_2 | 1.50 | Transmittance
12 | NCSR | Sunset on-line | EUSAAR_2 | 2.14 | Transmittance
13 | UHEL | Sunset on-line | EUSAAR_2 | 0.92 | Transmittance
Notes: In lack of certified reference material for atmospheric OC and EC, this ILCE made use of ambient (outdoor) PM aerosol collected with high-volume samplers on quartz fiber filters at 4 sites across Europe.
The homogeneity of these test samples was investigated by the JRC on distinct samples collected at each location. From each sample, ten subsamples of 1 cm² were taken along two perpendicular axes across the filter surface and analysed for their TC, OC and EC contents. The filter homogeneity was assessed as the standard deviation of the average of the 10 replicate analyses. This leads to an upper limit for the filter homogeneity since it includes the repeatability of the JRC laboratory (< 3 and 6% for TC and EC, respectively). The homogeneity is better than 6 and 11% for TC and EC/TC, respectively (Table 2). If sampling at each location occurred under repeatable conditions, it can be assumed that the test samples had similar homogeneities.
An aqueous solution of phthalic acid was also distributed to the participants to assess the uncertainty of the instrument calibration constant determination. The solution was prepared by dissolving a precisely known mass of pure phthalic acid (≥ 99.5%) in a precisely known volume of ultra-pure water (resistivity ≥18.2 mΩ cm).

## Slide 5
Repeatability and reproducibility standard deviations of the method for measurement of TC and EC/TC ratio
Evaluate the consistency of the dataset:
Mandel’s k parameter + Cochran’s test – within-laboratory consistency
Mandel’s h and G1-Grubbs’ test – between laboratory consistency
Identification of stragglers and outliers
Removal of outliers
Calculation of reproducibility and repeatability standard deviations of the method and their dependence from the mean value
Method performance
Notes: Ambient PM filter samples: In absence of certified reference material for atmospheric TC, OC and EC deposited on filters, the measurement method performance (par. 2.1) and laboratory performances (par. 2.2) were evaluated using atmospheric PM collected on filters as test samples.
In this report we focus on the TC loadings (in μg cm-2) and EC/TC ratios reported by each participants for each test sample. TC represents the most robust (and protocol-independent) output of TOA analyses, while EC/TC ratios are free from biases in the total carbon determination, and reflect possible differences in the OC/EC split determination among participants.

## Slide 6
Chart: Laboratory
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
4638 | 0.30511 | 0.68847 | 0.52878 | 2.105 | 0.1467 | 0.7015 | 0.92938 | 1.7119 | 0.41441 | 1.0211 | 0.4469
4642 | 0.98366 | 0.61192 | 1.8972 | 1.3324 | 0.091499 | 0.17724 | 0.44096 | 0.49365 | 0.73836 | 0.63547 | 1.6901
DEM2 | 3.0494 | 0.4641 | 0.28337 | 0.55887 | 0.39287 | 0.46686 | 0.24734 | 0.26528 | 0.33666 | 0.62812 | 0.28543
DEM3 | 1.3597 | 0.78371 | 0.42303 | 0.63929 | 1.0904 | 0.63034 | 1.8356 | 0.73694 | 0.69388 | 1.305 | 0.51583
T6694 | 0.38895 | 0.76663 | 0.66724 | 1.9909 | 0.64412 | 0.434 | 0.65283 | 0.55212 | 0.46874 | 1.8894 | 0.85345
T6696 | 0.79642 | 1.4643 | 0.74897 | 1.964 | 0.52869 | 0.96403 | 0.38175 | 1.3737 | 0.54708 | 0.42859 | 0.28041
A211 | 1.6581 | 1.2838 | 1.1852 | 0.60688 | 0.11732 | 0.32159 | 0.26535 | 0.6142 | 0.97116 | 1.3136
A212 | 1.1906 | 0.82537 | 1.2279 | 0.3171 | 0.59259 | 0.46685 | 1.9422 | 1.0362 | 0.42034 | 0.4567 | 1.2219
 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01
 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69
Chart: Laboratory
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
4638 | 0.95803 | -0.54617 | 0.99278 | -1.2882 | -1.7166 | 0.51485 | -0.78824 | 0.82487 | 1.1647 | 0.20223 | -0.47733
4642 | -1.395 | -0.15652 | 1.7129 | -0.27278 | -1.6248 | 0.024321 | -0.83092 | 0.34543 | 0.73101 | 0.57224 | 0.84189
DEM2 | 0.12002 | 0.53639 | 1.2286 | -2.2289 | -1.3457 | 0.067978 | -0.19569 | 0.6895 | 0.90644 | 0.26085 | -0.039556
DEM3 | -1.9181 | 0.88319 | 1.1949 | -1.4627 | -0.77593 | 0.29627 | -0.073514 | 0.62414 | 0.91197 | 0.10464 | 0.21503
T6694 | -1.4817 | 0.34577 | 1.8819 | -0.59329 | -1.2153 | -0.72308 | -0.43256 | 0.67955 | 0.33894 | 0.94449 | 0.36828
T6696 | -1.3192 | 0.085726 | 1.2397 | -1.3344 | -0.91453 | -0.29306 | -0.63728 | 1.61 | 0.66111 | 0.73384 | 0.19663
A211 | 0.078664 | -0.023388 | 0.63727 | -0.29761 | -0.47242 | -0.76409 | -0.33086 | 0.070971 | 0.25051 | 0.031917 | 0.81903
A212 | -0.055237 | -0.15898 | 1.3163 | -1.5181 | -0.98282 | -0.20011 | 0.5875 | 0.33924 | 0.27994 | -0.32719 | 0.71949
 | -2.22 | -2.22 | -2.22 | -2.22 | -2.22 | -2.22 | -2.22 | -2.22 | -2.22 | -2.22 | -2.22
 | -1.82 | -1.82 | -1.82 | -1.82 | -1.82 | -1.82 | -1.82 | -1.82 | -1.82 | -1.82 | -1.82
 | 1.82 | 1.82 | 1.82 | 1.82 | 1.82 | 1.82 | 1.82 | 1.82 | 1.82 | 1.82 | 1.82
 | 2.22 | 2.22 | 2.22 | 2.22 | 2.22 | 2.22 | 2.22 | 2.22 | 2.22 | 2.22 | 2.22
Mandel’s k parameter (repeatability)
Two outliers (lab/sample: 1/DEM2; 4/4638)
Eight stragglers (lab/sample: 3/4642;
4/T6694 and T6696; 7/DEM3 and A212; 8/4638; 10/T6694; 11/4642)
EUSAAR-2 performance for measurement of TC
Mandel’s h parameter (reproducibility)
One outlier (lab/sample: 4/DEM2)
Two stragglers (1/DEM3; 3/T6694)

## Slide 7
EUSAAR-2 performance for measurement of TC
test sample | general mean µgC / cm² | sr % | sR %
4638 | 21.5 | 2.8 | 6.4
4642 | 7.6 | 5.2 | 9.3
DEM2 | 13.5 | 1.9 | 7.0
DEM3 | 9.8 | 2.0 | 9.5
T6694 | 11.1 | 2.5 | 7.9
T6696 | 5.8 | 4.6 | 8.9
A211 | 4.9 | 5.3 | 10.4
A212 | 6.1 | 3.8 | 7.7
Chart: m
| 21.455582582349866 | 7.5969373786111998 | 13.489000323311737 | 9.7585948928566122 | 11.07983025783456 | 5.7953188223862426 | 4.940348525738723 | 6.1129929750245315
sR/m 2016 | 0.063642 | 0.093066 | 0.069763 | 0.095165 | 0.078567 | 0.089346 | 0.10441 | 0.076809
sr/m 2016 | 0.027876 | 0.052481 | 0.019019 | 0.019973 | 0.024958 | 0.046184 | 0.052511 | 0.038347
sR/m 2015 | 0.10456 | 0.089117 | 0.050103 | 0.067956 | 0.11628 | 0.11798 | 0.082061 | 0.067755
sr/m 2015 | 0.04964 | 0.068037 | 0.014643 | 0.015839 | 0.062394 | 0.050749 | 0.043841 | 0.04136
sR/m 2014 | 0.067823 | 0.085938 | 0.14181 | 0.07959 | 0.05034 | 0.057025
sr/m 2014 | 0.024798 | 0.062764 | 0.043461 | 0.046444 | 0.022823 | 0.014192
Method precision for TC measurement becomes exponentially poorer toward lower TC contents
i.e. 10 µgC/cm²

## Slide 8
EUSAAR-2 performance for measurement of EC/TC
Chart: Laboratory
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
4638 | 0.47551 | 0.31741 | 0.5212 | 0.84518 | 0.12234 | 0.51571 | 0.2822 | 2.7104 | 1.3143 | 0.45319 | 0.21828
4642 | 2.5692 | 0.13777 | 0.48067 | 1.7161 | 0.14783 | 0.6336 | 0.16226 | 0.54285 | 0.32187 | 0.48121 | 0.35305
DEM2 | 1.0831 | 0.51892 | 0.54003 | 0.54416 | 1.6633 | 1.2427 | 1.4559 | 0.78752 | 0.83469 | 0.46482 | 1.0031
DEM3 | 1.99 | 0.41102 | 1.1108 | 0.61996 | 0.4389 | 1.4984 | 0.92451 | 0.91723 | 0.86114 | 0.37766 | 0.48439
T6694 | 2.2571 | 0.77537 | 0.56698 | 0.87205 | 0.67978 | 1.3654 | 0.67305 | 0.25962 | 0.39561 | 1.1029 | 0.053106
T6696 | 2.4287 | 0.098041 | 0.20779 | 1.4779 | 0.11057 | 0.43278 | 0.21363 | 0.25964 | 0.3587 | 0.35094 | 1.5167
A211 | 1.4859 | 1.4929 | 0.40951 | 0.59526 | 0.67357 | 0.58576 | 0.67127 | 1.5297 | 1.0536 | 0.52572 | 1.0333
A212 | 1.0752 | 0.43325 | 0.80708 | 0.66339 | 1.39 | 0.39643 | 1.6271 | 1.1054 | 1.228 | 0.56442 | 0.88305
 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01 | 2.01
 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69 | 1.69
Chart: Laboratory
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
4638 | 1.6744 | -0.607 | -0.028264 | 0.69391 | -0.37709 | 0.81757 | -0.26382 | 0.28144 | -0.30249 | -0.3751 | -2.2703
4642 | 2.045 | -1.1981 | 0.070279 | 0.1428 | -0.54313 | 0.83577 | 0.61764 | -0.13686 | -0.73626 | 0.058387 | -1.5549
DEM2 | 2.2562 | -0.99151 | 0.21896 | 0.74264 | -0.40654 | 0.78519 | -0.15619 | -0.24337 | -0.51053 | -0.22907 | -1.4658
DEM3 | 0.68853 | -0.90766 | -0.20163 | -0.11559 | -0.12002 | 2.5832 | 0.27531 | -0.40453 | -0.3889 | -0.16688 | -1.2418
T6694 | 1.8915 | -1.093 | 0.42348 | 0.34935 | -0.64443 | 1.1531 | 0.57318 | -0.26158 | -0.46299 | -0.59363 | -1.4892
T6696 | 1.0087 | -1.4722 | 0.51406 | 0.6781 | -0.5396 | 0.56481 | 1.5457 | -0.23593 | -0.31257 | -0.65984 | -1.5821
A211 | 0.20552 | -0.13762 | 1.5157 | 0.50367 | -0.6021 | 0.73681 | 0.71648 | 0.38855 | -0.62841 | -0.4433 | -2.2553
A212 | 1.2705 | -0.56779 | 1.2071 | 0.78768 | -0.77064 | 0.51212 | -0.15081 | 0.097357 | -0.19073 | -0.020388 | -2.1743




Mandel’s k parameter (repeatability)
Four outliers (lab/sample: 1/4642, T6694 and T6696; 8/4638)
Two stragglers (lab/sample: 1/DEM3; 4/4642
Mandel’s h parameter (reproducibility)
Four outliers (lab/sample: 1/DEM2; 6/DEM3; 11/4638 and A211)
Three straggler (1/4642, T6694; 11/A212)
The recurrence of stragglers and outliers from a single laboratory most probably suggests unsatisfactory precision for EC/TC ratio determination as compared to the other participants

## Slide 9
EUSAAR-2 performance for measurement of EC/TC
test sample | general mean ratio | sr % | sR %
4638 | 0.19 | 4.2 | 11.9
4642 | 0.16 | 4.7 | 11.3
DEM2 | 0.15 | 5.2 | 15.8
DEM3 | 0.10 | 4.9 | 16.6
T6694 | 0.15 | 3.8 | 17.2
T6696 | 0.12 | 6.1 | 19.2
A211 | 0.13 | 7.0 | 24.4
A212 | 0.12 | 6.0 | 19.3
Chart: m
| 0.18711649948980313 | 0.15573295684066424 | 0.14731428075841926 | 9.9985573580120043E-2 | 0.15435691425110182 | 0.12301794487612402 | 0.12896099246430515 | 0.12146344509600027
sR/m 2016 | 0.11947 | 0.11265 | 0.15821 | 0.16585 | 0.17206 | 0.19189 | 0.24403 | 0.19299
sr/m 2016 | 0.041541 | 0.046502 | 0.052152 | 0.048788 | 0.037804 | 0.060686 | 0.069717 | 0.060054
sR/m 2015 | 0.26246 | 0.18737 | 0.12898 | 0.17336 | 0.24173 | 0.23769 | 0.21167 | 0.17338
sr/m 2015 | 0.088441 | 0.087462 | 0.059978 | 0.062109 | 0.11075 | 0.092025 | 0.093234 | 0.072486
sR/m 2014 | 0.18197 | 0.6586 | 0.30651 | 0.21749 | 0.21086 | 0.22298
sr/m 2014 | 0.035599 | 0.15982 | 0.12131 | 0.098357 | 0.031448 | 0.02472
Method precision for EC/TC ratio measurement becomes exponentially poorer toward lower EC/TC ratio. i.e. 0.05

## Slide 10
Laboratory performance for measurement of TC:
z-scores
z-scores describe the laboratory bias compared to the assigned value associated with its standard deviation
z = (xi-X)/ σ
Assigned value and standard deviation: “Consensus value from participants”, i.e. robust average and standard deviation
TC |  | 4638 | 4642 | DEM2 | DEM3 | T6694 | T6696 | A211 | A212
assigned value | μg/cm2 | 21.7 | 7.6 | 13.6 | 9.8 | 11.1 | 5.8 | 4.9 | 6.1
standard deviation | μg/cm2 | 1.4 | 0.7 | 0.7 | 0.9 | 0.9 | 0.5 | 0.5 | 0.4
1σ | % | 6.3 | 9.0 | 4.9 | 9.1 | 7.9 | 9.1 | 10.3 | 6.1
2σ | % | 13 | 18 | 10 | 18 | 16 | 18 | 21 | 12
3σ | % | 19 | 27 | 15 | 27 | 24 | 27 | 31 | 18

## Slide 11
Laboratory performance for measurement of TC:
z-scores
Chart: laboratory
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
4638 | 0.69335 | -0.68656 | 0.72523 | -1.3673 | -1.7603 | 0.28679 | -0.90862 | 0.57119 | 0.88292 | 0 | -0.6234
4642 | -1.2633 | -0.13501 | 1.5682 | -0.24092 | -1.4727 | 0.02975 | -0.74942 | 0.3223 | 0.67358 | 0.52893 | 0.7746
DEM2 | -0.067758 | 0.47356 | 1.3736 | -3.1216 | -1.9733 | -0.13542 | -0.47823 | 0.67263 | 0.95468 | 0.11534 | -0.27523
DEM3 | -2.0513 | 0.83006 | 1.1507 | -1.5828 | -0.87647 | 0.22637 | -0.15399 | 0.5636 | 0.85967 | 0.029258 | 0.1428
T6694 | -1.3981 | 0.33586 | 1.7934 | -0.55516 | -1.1454 | -0.67831 | -0.40265 | 0.65256 | 0.32938 | 0.90395 | 0.35721
T6696 | -1.1656 | 0.073306 | 1.0909 | -1.179 | -0.80876 | -0.26072 | -0.56427 | 1.4175 | 0.5807 | 0.64484 | 0.1711
A211 | 0.15674 | -0.046312 | 1.2682 | -0.59194 | -0.93975 | -1.5201 | -0.6581 | 0.14144 | 0.49866 | 0.063731 | 1.6299
A212 | -0.099445 | -0.25096 | 1.9036 | -2.2359 | -1.4541 | -0.31103 | 0.83922 | 0.47665 | 0.39006 | -0.49661 | 1.032
 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
One outlier (lab/sample; 4/DEM2) and two stragglers (lab/sample: 1/DEM3; 4/A212)
80% of all entries is within 10% from the assigned value and 95% within 15%
A few participants show the systematic tendency of underestimating or overestimating the assigned TC concentrations

## Slide 12
EC/TC |  | 4638 | 4642 | DEM2 | DEM3 | T6694 | T6696 | A211 | A212
assigned value | ratio | 0.18 | 0.16 | 0.14 | 0.10 | 0.15 | 0.12 | 0.13 | 0.12
standard deviation | ratio | 0.01 | 0.02 | 0.02 | 0.01 | 0.03 | 0.03 | 0.02 | 0.02
1σ | % | 6 | 12 | 12 | 11 | 17 | 22 | 19 | 17
2σ | % | 12 | 23 | 24 | 22 | 34 | 43 | 38 | 33
3σ | % | 18 | 35 | 36 | 33 | 51 | 65 | 57 | 50
Chart: laboratory
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
4638 | 3.7203 | -0.85691 | 0.30423 | 1.7532 | -0.39563 | 2.0013 | -0.16837 | 0.9256 | -0.24596 | -0.39164 | -4.194
4642 | 2.7732 | -1.1941 | 0.35752 | 0.44623 | -0.39285 | 1.2939 | 1.0271 | 0.10412 | -0.6291 | 0.34297 | -1.6305
DEM2 | 3.0569 | -1.1298 | 0.43065 | 1.1057 | -0.37568 | 1.1606 | -0.052963 | -0.16535 | -0.50974 | -0.14691 | -1.7411
DEM3 | 2.4241 | -1.8594 | 0.035232 | 0.26613 | 0.25424 | 7.5085 | 1.3151 | -0.50925 | -0.46731 | 0.1285 | -2.7562
T6694 | 1.9399 | -1.0099 | 0.48896 | 0.41569 | -0.56656 | 1.2101 | 0.63692 | -0.18815 | -0.38722 | -0.51634 | -1.4015
T6696 | 1.0422 | -1.1851 | 0.59809 | 0.74535 | -0.34784 | 0.64365 | 1.5243 | -0.075224 | -0.14402 | -0.45579 | -1.2837
A211 | 0.13044 | -0.29318 | 1.7478 | 0.49852 | -0.86659 | 0.78635 | 0.76124 | 0.3564 | -0.89908 | -0.67054 | -2.9075
A212 | 1.3358 | -0.72266 | 1.2648 | 0.79516 | -0.94981 | 0.48659 | -0.25574 | 0.022152 | -0.30044 | -0.1097 | -2.5216
 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
Laboratory performance for measurement of EC/TC:
z-scores

## Slide 13
Laboratory performance for measurement of EC/TC:
z-scores
Chart: laboratory
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
4638 | 3.7203 | -0.85691 | 0.30423 | 1.7532 | -0.39563 | 2.0013 | -0.16837 | 0.9256 | -0.24596 | -0.39164 | -4.194
4642 | 2.7732 | -1.1941 | 0.35752 | 0.44623 | -0.39285 | 1.2939 | 1.0271 | 0.10412 | -0.6291 | 0.34297 | -1.6305
DEM2 | 3.0569 | -1.1298 | 0.43065 | 1.1057 | -0.37568 | 1.1606 | -0.052963 | -0.16535 | -0.50974 | -0.14691 | -1.7411
DEM3 | 2.4241 | -1.8594 | 0.035232 | 0.26613 | 0.25424 | 7.5085 | 1.3151 | -0.50925 | -0.46731 | 0.1285 | -2.7562
T6694 | 1.9399 | -1.0099 | 0.48896 | 0.41569 | -0.56656 | 1.2101 | 0.63692 | -0.18815 | -0.38722 | -0.51634 | -1.4015
T6696 | 1.0422 | -1.1851 | 0.59809 | 0.74535 | -0.34784 | 0.64365 | 1.5243 | -0.075224 | -0.14402 | -0.45579 | -1.2837
A211 | 0.13044 | -0.29318 | 1.7478 | 0.49852 | -0.86659 | 0.78635 | 0.76124 | 0.3564 | -0.89908 | -0.67054 | -2.9075
A212 | 1.3358 | -0.72266 | 1.2648 | 0.79516 | -0.94981 | 0.48659 | -0.25574 | 0.022152 | -0.30044 | -0.1097 | -2.5216
 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3 | -3
 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2 | -2
 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
Four outliers (lab/sample: 1/4638 and DEM2; 6/DEM3; 11/4638) and six stragglers (1/4642 and DEM3; 6/4638; 11/DEM3, A211, A212) all produced by three participants
Only 50% of all entries is within 10% of the assigned value and 85% is within the 25% of the assigned EC/TC ratio
The majority of participants show the systematic tendency of underestimating or overestimating the assigned EC/TC ratio

## Slide 14
The measurement method repeatability and reproducibility
for TC ranged from 2% to 5% and from 6% to 10% respectively.
for the EC/TC ratio from 4% to 7 and from 11% to 24% respectively.
combining the repeatability and reproducibility relative standard deviation for the EUSAAR-2 protocol obtained during the 2014, 2015 and the present ILCE, we observed that the method precision becomes exponentially poorer toward lower TC contents i.e. 10 µgC / cm² and EC/TC ratio. i.e. 0.05
Conclusions-1

## Slide 15
Laboratory performances based on z-scores:
- For TC loadings, only one outlier and two stragglers were identified;
80% of all entries is within 10% from the assigned TC concentration value and 95% within 15%.
A few participants show the systematic (i.e. for all test samples) tendency of underestimating or overestimating the assigned TC concentrations.
Corrective action: more accurate determination of the calibration constant (e.g. implementing where possible CO2 calibration)
- For EC/TC ratios, all ten outliers and stragglers were produced by three participants. Only 50% of all entries is within 10% of the assigned value and 85% is within the 25% of the assigned EC/TC ratio.
The majority of participants show the systematic tendency of underestimating or overestimating the assigned EC/TC ratio.
Corrective action: a more solid and stable in time instrument set-up in terms of i) laser stability; ii) FID response in He and He/O2 phases; iii) temperature calibration and iv) transit
Conclusions-2

## Slide 16
Conclusions-3
Participation to inter-laboratory comparison is beneficial (if not compulsory) to users and to the scientific community
ECAC is organizing a second ILCE in 2016:
Filters: end August Analysis: mid October Deadline for application: 1 st July

## Slide 17
Reference Material
In April 2015 JRC started to design and test a way to produce a reference material for measurement of atmospheric particulate organic and elemental carbon deposited on filters
Several organic substances and elemental carbon forms were investigated for their ability of making pyrolysis in a inert atmosphere (the organic compounds) and decomposing in an oxidative atmosphere at temperature lower than 850 degree Celsius (the elemental carbon forms)
The species simulating at best the organic and elemental carbon aerosol fractions were selected to generate a suspension
Ink-jet printing on the filters, using the suspension as 'ink', has been shown to be the most convenient approach
Your results would contribute to the homogeneity and transport stability study